## Slide 1: The Moon as Possible Calibration Reference for Microwave Radiometers
M. J. Burgdorf, Universität Hamburg, Germany, S. A. Bühler, Universität Hamburg, Germany, I. Hans, Universität Hamburg, Germany, T. Lang, Universität Hamburg, Germany, S. Michel, Universität Hamburg, Germany, O. Lemke, Universität Hamburg, Germany
Summary: Instruments on satellites for Earth observation on polar orbits usually employ a two-point calibration technique, in which deep space and an on-board calibration target provide two reference flux levels. As the direction of the deep space view is close to the celestial equator, the Moon moves sometimes through the field of view and introduces an unwelcome additional signal. One can take advantage of this intrusion, however, by using the Moon as a third flux standard for stability analysis of microwave sensors. This method is relevant for FIDUCEO, a project producing fundamental climate data records with traceable estimates of stability and thereby enhanced credibility.
Position of the Moon When Seen by the DSV
A cross-scanning instrument takes data along a great circle perpendicular to its flight direction. The DSV position is then usually within 20° of the orbital axis of the spacecraft. It describes a circle in the sky, whose center is 9° away from the celestial equator. The Moon crosses this circle several times per year and is seen during several consecutive orbits.
Instr./Sat. | Cone angle θ | Intrusions/year | Eq. x-ing | -Sat-
MHS/MetOp-B | 32.4° - 37.2° | 7 groups of orbits | 21:30 | 110° - 145°
AMSU-A/N-15 | 13.4° - 26.6° | 5 groups of orbits | 17:30 | 168° - 180°
AMSU-B/N-16 | 32.4° - 37.2° | 4 groups of orbits | 16:30 | 144° - 174°
Strength of the additional signal of the Moon in the DSV
When using the Moon as a flux reference, it is important to have it close to the center of the DSV. This can be achieved with dedicated spacecraft maneuvers. The time difference Δt between minimum angle and maximum signal allows to estimate the tangential pointing error:
Δt / 100min × 360° × sin(θ/2)
The radial pointing error can be estimated from the maximal signals in consecutive orbits.
Obviously the brightness temperature of the Moon depends on its phase φ, and because its absolute radiance is not well known at all frequencies, its best use is for checking the stability of the calibration with observations at fixed φ.
Moon and Antenna Pattern
The Moon can be present at the same time in all four DSVs of MHS or AMSU-B (left). The comparison of the light curve (red) with a Gaussian (blue) betrays an asymmetric antenna pattern (right).
Object | TB, 24 GHz [K] | TB, 183 GHz [K] | Diameter arcmin | Φ24 Hz | Φ183 GHz
OBCT | 280 | 280 | - | 1 | 1
CMB | 2.725 | 2.725 | - | 7.9×10-3 | 1.3×10-3
Luna | 258 | 258 | 30 | 2.1×10-2 | 1.9×10-1
Jupiter | 130 | 150 | 0.8 | 7.6×10-6 | 7.8×-5
Crab | S = 400 Jy | S = 290 Jy | 10 | 3.2×10-5 | 2.3×10-6
Brightness of other foreground objects
An analysis of possible error sources in the cold space radiance requires an estimate of the foreground emission. Jupiter, for example, would add about 20 counts to the cold calibration level of some 10,000 counts – barely detectable.
Frequency of Intrusions of the Moon in the DSV
The angle of the deep space view with the orbital axis of the satellite θ/2 determines ultimately the frequency of the Moon glint. But also the equator crossing time of the satellite is relevant, because it fixes the direction of the deep space view in the sky. The Moon can only have a narrow range of phase angle φ when seen by the instrument and must not be too far away from the celestial equator at this position in its orbit.