## Slide 1: Calibration of relative humidity instruments in low pressure, low temperature, CO2 environment
M. Genzer, J. Polkko, T. Nikkanen, M. Hieta, A.-M. Harri

## Slide 2
Pressure in order of ~ 6 - 10 hPa
Night temperatures down to -70°C and lower
  Pw < 1 Pa even in saturated conditions
96% CO2 atmosphere
Mars
Humicap® based relative humidity device
NASA/JPL-Caltech/MSSS
Temperature / °C
Capacitance
Dry points
Saturation points
Change of Humicap® capacitance as function of temperature and relative humidity. The capacitance affected also by CO2.

## Slide 3
Dry conditions
EASY
Saturated conditions HARD
Minimum requirements for calibration conditions
-70°C …+25°C
-70°C …+25°C
Constant gas flow through the chamber for uniform conditions

## Slide 4
The valve and pump system enables different routes for the gas (air or CO2) to the pressure chamber inside the temperature cabinet. Also the pressure inside the vessel is controlled by opening and closing the valves. Measurements in ambient conditions, vacuum, low-pressure air and low-pressure CO2 (dry or wet) are all possible with the same setup.
Calibration setup at FMI

## Slide 5
For measuring dry points in cold, low-pressure CO2 environment, dry CO2 is routed to the pressure chamber with a constant flow that can be adjusted with the valves until the desired reference pressure is achieved.
Measurements are made in several temperature points ranging from room temperature down to -70°C. For comparative measurements in vacuum the CO2 is switched off and pumping efficiency increased.
Dry point measurements

## Slide 6: CO2 effect on Humicap®
Chart: REMS-H REF difference in capacitance between vacuum and ~8 hPa dry CO2
| -10.2 | -30.0 | -49.8 | -69.5
Humi 1 diff | -0.0069366 | -0.020243 | -0.045744 | -0.084124
Humi 2 diff | -0.0066158 | -0.020047 | -0.045073 | -0.083453
Humi 3 diff | -0.0070625 | -0.020394 | -0.045703 | -0.084105
average diff | -0.0068716 | -0.020228 | -0.045507 | -0.083894
Chart: DREAMS-H ‘d’ Difference in capacitance between vacuum and ~8hPa dry CO2
| -9.9 | -29.9 | -49.5 | -69.4
pF difference average | -0.0072184 | -0.018762 | -0.044218 | -0.07549
Humi 1 pF difference | -0.00707 | -0.018668 | -0.044234 | -0.075564
Humi 2 pF difference | -0.0073974 | -0.01915 | -0.045233 | -0.077056
Humi 3 pF difference | -0.0071878 | -0.018467 | -0.043188 | -0.073848
The CO2 effect is larger in higher pressures and cold temperatures. Taking into account also the dynamic range of the Humicap®‚ in Martian pressure range the effect is negligible in temperatures > ~ -10°C.

## Slide 7
To achieve saturated conditions, CO2 is routed via a vessel filled with water. Moist CO2 is then routed with a constant flow to the pressure chamber. After some time the chamber reaches stable, saturated conditions where the chamber walls are covered with a thin ice layer.
Measurements in CO2 saturated with water vapor

## Slide 8
Switch from dry to wet gas
Humicap® capacitance reaction to wet gas
Chart: -60°C, 9.5 hPa, CO2+H2O
Chart: -60°C -> -70°C, 9.5 hPa, CO2+H2O
4 8 12 16 20 24 28 32 36 40 44 48 52 56 60 64 68 72
…
After switching from dry to wet CO2, the conditions inside the chamber stabilize in approx. 24h, which can be seen from the steady measurements of Humicap® capacitance. Changing the temperature of the chamber to even colder temperature results in Humicap® setting itself to a new level.
Temp change -60°C -> -70°C

## Slide 9
Temp switch -70°C -> -60°C
Chart: -70C -> -60°C, 9.5 hPa, CO2+H2O
When the temperature is switched back to original -60°C, the Humicap® capacitance also returns to the original level, which means that the conditions inside the chamber are stable.
0 2 4 6 8 10 12 14 16 18 20 22 24 26

## Slide 10
How about switching back to using only dry CO2?

## Slide 11
Switch to dry CO2
Chart: -60°C, 9.5 hPa, switch to dry CO2
0 2 4 6 8 10 12 14 16 18 20 22 24 26 28
The layer of ice inside the chamber is enough to keep up the saturated conditions even after the gas flow has been switched back to dry CO2.

## Slide 12: Conclusions and future
FMI humidity calibration chamber can achieve stable dry and saturation points in
  CO2 environment (or air)
  Temperatures down to -70°C
  Martian pressure range and lower (down to vacuum)
    Higher pressures are possible too, but then the lower temperature in CO2 is constrained
Future development:
  Comparison measurements with other facilities
  Intermediate humidity points between dry and saturation ..?