## Slide 1
Study on the correction of total suspended particles monitoring results based on machine learning method
Qiaofeng Guo, Zhen Cheng
Shanghai Jiao Tong University

## Slide 2
CONTENTS
Research background
1
Methods and data sources
2
Results of ML models
3
4
Analysis of influencing factors
5
Conclusion
2

## Slide 3
Part 1
Reseach background
3

## Slide 4
(Bulletin of Shanghai ecological environment, 2018)
Research background
—Road dust pollution
From 2018, Shanghai will no longer carry out conventional manual monitoring of regional dust reduction, but instead carry out online monitoring of road dust.
However, there is a big error between the results of the commonly used on-line measurement of TSP light scattering instrument and the results of the weighing method。
(Dust data online monitoring platform of Shanghai Eureka Environmental Protection Hi-tech., Ltd)

## Slide 5
Machine learning correction / prediction models for fine particles and gaseous pollutants have been studied.
(Zimmerman et al., AMT, 2018)
(Xiao et al., ES&T, 2018)
Research background
—Machine Learning correction

## Slide 6
Research Contents
Using machine learning model to correct the TSP measurement result of light scattering monitoring instrument
Using the interpretable tools in machine learning, this paper analyzes the influence of various pollutants and environmental temperature and humidity on the instrument monitoring

## Slide 7
Part 2
Methods and data sources

## Slide 8
TSP Measurement
Casella dust monitor
CLE-712
—light scattering and gravimetric methods
Parameter | Method | Instrument | Model | Factory number
TSP | GB/T 15432-1995 | Balance | MS105DU | B524045289
TSP | GB/T 15432-1995 | Chamber | LF-3000 | SHNX-2018-B0913
TSP | GB/T 15432-1995 | TSP Sampler | 2030 | M03358884、M03357505
TSP | GB/T 15432-1995 | Flow calibrator | 8040 | 2L01021892
Temperature and humidity | GB/T 18204.1-2013 | Hygrothermograph | 8809 | 9152104
National standard method
GB/T15432-1995

## Slide 9
Method
—Machine learning algorithm
ML algorithm | Brief introduction
SVM | Good adaptability to high dimensional space and memory saving
RF | Strong anti over fitting ability
GBRT | Strong ability to handle different types of data
ANN | Good at analyzing large amount of data
SVM
RF
GBRT
ANN
TSP Gravimetric results
Casella raw readings
PM concentrations (PM2.5、PM10)
Gaseous pollutants concentrations
(SO2、NO2、CO、O3)
Tem, RH
INPUT
Monitoring time (hour)
OUTPUT
ML Models

## Slide 10
Using the importance analysis tool of random forest to rank the importance of each input feature.
(Hu et al., ES&T, 2017)
Method
—Interpretable tools (RF importance analysis)

## Slide 11
(Sayegh et al., AE, 2016)
Partial dependence plots blurs the influence of other features and shows the marginal effect of target features on the results.
Method
—Interpretable tools (Partial dependence plot)

## Slide 12
Data source
1114 hourly records were selected in 17 TSP monitoring sites and 7 national monitoring stations in urban Shanghai.

## Slide 13
Part 3
Results of ML models

## Slide 14
Comparison of original data and training effect
—CaoXi Road site
Notes: EP1

## Slide 15
Comparison of original data and training effect
—CaoXi Road site
Notes: EP1

## Slide 16
—CaoXi Road site
Comparison of original data and training effect
Notes: EP1

## Slide 17
Comparison of model test results

## Slide 18
Four models have greatly improved the correlation between the two methods, the slope range is 0.91-1.11, and the correlation coefficient is above 0.75
SVM
RF
GBRT
ANN
RMSE=49.01
MAE=24.19
RMSE=49.53
MAE=24.71
RMSE=53.49
MAE=30.15
RMSE=52.19
MAE=29.18
Comparison of model test results

## Slide 19
Part 4
Analysis of influencing factors

## Slide 20
Fearture Importance — RF feature importance and ablation study of three models
PM2.5, RH and TSP are the three features that have the greatest influence on TSP measurement of light scattering instrument

## Slide 21
Correlation analysis between features
The result shows that the relationship between PM2.5 and PM10 is very close, which has a correlation coefficient of 7.3. RH has a negative correlation with most features (O3, PM2.5, PM10, TSP, Tem).

## Slide 22
Correlation analysis of feature to target value — PDP

## Slide 23
Two-way PDP charts

## Slide 24
Two-way PDP charts

## Slide 25
Comparison with previous studies — scattering hygroscopic growth curves
The f(RH) values of PM2.5 from previous studies are about 1 to 3 times that of TSP in this study, indicating that the scattering hygroscopic efficiency of PM2.5 is much higher than that of TSP.

## Slide 26
Part 5
Conclusion

## Slide 27
Conclusion
All four machine learning models greatly improved the correlation of the TSP concentrations reported by the Casella with filter-based measurements.
PM2.5 has the largest aerosol scattering efficiency, so it has a great influence on the measurement value of light scattering instrument.
Humidity has a significant impact on the monitoring of particulate matter at any concentration, and it is necessary to correct the high humidity data.

## Slide 28
应用前景
Related Publications
[1] Qiaofeng Guo, Zhu Zhu, Zhen Cheng*, Shuhong Xu, Xiaoliang Wang, Yusen Duan. Correction of Light Scattering-based Total Suspended Particulate Measurements through Machine Learning[J]. Atmosphere, 2020.

## Slide 29
THANKS!
Thanks for data support from Shanghai Eureka Environmental Protection Hi-tech company and Shanghai Environmental Monitoring Center