## Slide 1: Correction of calibration errors for air humidity measurements Fyodor Tatarinov, Jonathan Muller, Eyal Rotenberg and Dan Yakir Weizmann Institute of Science, Rehovot, Israel

## Slide 2: Motivation
Infrared gaze analyzers (IRGAs) used in Eddy Covariance (EC) measurements can have a drift of water concentration (Cw) measurement error with time, what requires regular calibration (from our experience once per 1-2 weeks). Accurate correction of errors between calibrations is not usually performed.
Capacitive humidity sensors (CS, like, e.g., Rotronic or Vaisala) have slow drift, but are less precise.
Accurate RH measurement is necessary, e.g., for getting canopy humidity profiles or fixing dew deposition events

## Slide 3: Yatir FLUXNET site
Northern edge of Negev desert, Israel, 31O20'N, 35O03'E, about 650 m a.s.l.
Aleppo pine forest planted in the mid-1960s, canopy height ~10 m
Summer drought period from May to October, winter period with moderate precipitation (280 mm on average).
Eddy covariance measurements since 2000

## Slide 4: Air humidity and related measurements in Yatir
* June 04 moved to 5 m
Device | Level, m | Variable, units | Quality | Calibration
LI-7000 | 19 | Cw, mmolmol-1 | Rather accurate, drift within days/weeks | Once a week to once a month
LI-6262 | 1, 9, 13, 19 | Cw, mmolmol-1 | Less accurate, drift within days/weeks | Once a week to once a month
Rotronic | 1, 15* | RH, % T, oC | Slow drift (years), not enough accurate for high and low RH | No calibration
Thermocouples | 1, 7.5, 9, 13, 15, 19 | T, oC |  | No calibration
*Further we assume that difference in Cw at 19 and 15 m is negligible, which allow us to compare Rotronic 15m and IRGA 19 m data

## Slide 5: Basic assumptions.
The aim of the project
To calculate the most accurate air humidity estimates above canopy for all ½-hourly records when such data are available.
To correct LI-6262 readings by comparison of its 19 m values with corrected simultaneous 19 m measurements and apply correction to other levels
Data of LI-7000 are accurate after calibration
Rotronic is drifting much slower than LI-7000 and LI-6262
Rotronic is not relevant for low and high RH
The difference between corrected absolute humidity (Cw) at 15 m (Rotronic) and 19 m (LI-7000 and LI-6262) can be neglected.

## Slide 6: Schematic of the problem
Chart: Absolute humidity, relative units
Theoretically we could expect gradual accumulation of error after IRGA calibration (which we could for simplicity assume to be linear), and relatively stable CS (Rotronic) error – something like on this figure. In this case we can simply interpolate correction factor from 0 after calibration to max before next calibration
Cumulated error (red line)
Calibration events
Cumulated error (yellow line)

## Slide 7: Therefore we developed an algorithm for double-correction of air humidity readings
Non-linear or even non-monotonous changes of IRGA error;
Too big difference between IRGA and CS in some periods;
Missing IRGA data before calibration, which do not allow to evaluate inter-calibration drift
Problems with field calibration (too big variation of IRGA readings during calibration)
E.t.c.
However real situation can be more complicated… In particular, in some periods it can happen:

## Slide 8: Order of calculations
Converting Rotronic 15 m relative humidity (RH, %) into absolute humidity (Cw, in mmolmol-1)
Calculating correction factors for Rotronic for each evenings of LI-7000 calibration days :
,
where Cw was averaged for first 3 h after sunset, excluding records, when RHRotronic<25% or RHRotronic>85%. Calibration events with |dCwCS|>0.8 mmolmol-1 were excluded.
dCwCS was interpolated between calibration events and corrected CwCS was calculated as CwCS,corr = CwCS + dCwCS
CwIRGA,corr was used to get the Cw correction factors for for both IRGAs, LI-7000 and LI-6262 at 19 m, on daily basis for 3 h after sunset (when RHRotronic is 25-85%):
dCwIRGA =CwIRGA - CwCS,corr
Then dCwIRGA was interpolated
Corrected IRGA values are calculated as:
  CwIRGA,corr = CwIRGA - dCwIRGA

## Slide 9: Results – correction factor of Rotronic
Chart: Corr_factor_Rotronic
Rotronic air humidity is really rather stable, although a small trend can be detected

## Slide 10: Results – correction factors for IRGA
LI-Cor, both LI-7000 and especially LI-6262, showed considerably higher instability than Rotronic. However, LI-7000 had the error of similar magnitude as Rotronic during rather long periods, e.g., 07.2017-02.2018, 09.2019-01.2020
Chart: Correction factor, mmol mol-1

## Slide 11: Results – corrected Cw values
Chart: Absolute humidity, mmol mol-1

## Slide 12: Conclusions
IRGAs humidity readings are not stable, regular calibration is needed
A method of correction of air humidity measurement errors was developed.
Apparently we got the best possible evaluation of air humidity using equipment available.
After correction the accuracy of IRGA Cw readings considerably improved. Median absolute deviation between Rotronic and IRGA decreased from 0.57 down to 0.26 mmolmol-1 for LI-7000 and from 0.87 to 0.30 mmolmol-1 for LI-6262