## Slide 2: Overview
Intro
Factors affecting the accuracy of PARIO
ISP+
Test with real data

## Slide 3: Particle size distribution determines soil texture
Example:
20 % sand 60 % silt 20 % clay  „silt loam“

## Slide 4: Cumulative particle-size distribution curve is obtained by sieving and sedimentation
silt
sand
gravel
clay

## Slide 5: Three alternative methods for sedimentation analysis
HYDROMETER
Köhn, 1931
Casagrande, 1934
Durner et al., 2017
Sedimentation of particles: pedological reference method for particle size distribution (PSD)
Particles settle with size-dependent velocity in a sedimentation cylinder.
Sinking fronts are formed and the particle size distribution can be determined traditionally by taking an aliquot of the suspension (PIPETTE method) or by measuring the buoyancy of a body (HYDROMETER method).
A recent variant is the integral suspension pressure method (ISP), in which the suspension pressure is measured continuously at one depth and the particle size distribution is determined by an inverse simulation of the change in pressure signal over time.

## Slide 6: Overview
Intro
Factors affecting the accuracy of PARIO
ISP+
Test with real data

## Slide 7: Accuracy of PARIO with ISP
Practical tests:
accuracy of the ISP is less than theoretically expected ( ±1.5 %)
measurement time exceeds it’s theoretical expectation (6 h).

## Slide 8: ISP classic: Low accuracy for clay fraction

## Slide 9: Reason: Variability of initial pressure decrease
Chart: Runtime [s]
Chart: Runtime [s]

## Slide 10: Causes and remedies
   Mixing (overhead shaking vs. stirring)
   Temperature effects
   Errors in soil mass (air dry vs. oven-dry )
   Organic matter destruction
   Mass of dispersant

## Slide 11
1,5 h am präzisesten, Abweichung zum Erwartungswert kann noch verringert werden
Schluffgehalte im Verhältnis zu Tongehalten nur ungenau bestimmt
Oxidation organischer Substanz beschleunigen
Auf Ablass nicht komplett verzichten, aber auf zweistufigen Ablass
Konstante Temperatur während Messung gewährleisten
Influence of mixing
   Mixing wets all the material above the supension level and causes retarded flow on walls into the suspension in the early measurement phase

## Slide 12: Pressure change by water flow from cylinder wall
Chart: Runtime [s]
| 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140
Pressure | 0 | 0.11 | 0.18 | 0.23 | 0.29 | 0.34 | 0.36 | 0.41 | 0.43 | 0.46 | 0.48
 | 0 | 0.44 | 0.59 | 0.71 | 0.76 | 0.82 | 0.87 | 0.89 | 0.94 | 0.95 | 0.98 | 1.01
 | 0 | 0.13 | 0.23 | 0.31 | 0.32 | 0.36 | 0.42 | 0.44 | 0.47 | 0.5 | 0.52
 | 0 | 0.53 | 0.71 | 0.82 | 0.89 | 0.96 | 1 | 1.06 | 1.1 | 1.12 | 1.12
 | 0 | 0.11 | 0.38 | 0.49 | 0.59 | 0.65 | 0.68 | 0.71 | 0.72 | 0.76 | 0.78
 | 0 | 0.25 | 0.41 | 0.49 | 0.57 | 0.64 | 0.65 | 0.68 | 0.68 | 0.71 | 0.73
 | 0 | 0.46 | 0.64 | 0.77 | 0.84 | 0.92 | 0.96 | 1.01 | 1.04 | 1.08 | 1.1
 | 0 | 0.28 | 0.43 | 0.48 | 0.55 | 0.57 | 0.62 | 0.66 | 0.67 | 0.69 | 0.73
 | 0 | 0.37 | 0.56 | 0.65 | 0.74 | 0.76 | 0.82 | 0.85 | 0.85 | 0.91 | 0.94
 | 0 | 0.34 | 0.54 | 0.68 | 0.75 | 0.82 | 0.89 | 0.92 | 0.94 | 0.97 | 0.99 | 1.01
Mittelwert Druckdifferenz | 0 | 0.32333 | 0.49889 | 0.6 | 0.66778 | 0.72222 | 0.76778 | 0.80222 | 0.82333 | 0.85444 | 0.87667
Chart: Runtime [s]
| 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140
 | 0 | 0.017 | 0.022 | 0.037 | 0.033 | 0.021 | 0.017 | 0.034 | 0.052 | 0.06 | 0.047 | 0.046
 | 0 | 0.009 | 0.019 | -0.003 | 0.035 | 0.044 | 0.023 | 0.032 | 0.04 | 0.04 | 0.038
 | 0 | -0.012 | -0.003 | 0.015 | 0.032 | 0.031 | 0.028 | 0.047 | 0.027 | 0.024 | 0.042 | 0.022
 | 0 | 0.011 | 0.029 | 0.038 | 0.055 | 0.034 | 0.053 | 0.121 | 0.12 | 0.129 | 0.159
 | 0 | 0.011 | 0.011 | 0.019 | 0.028 | 0.057 | 0.017 | 0.034 | 0.044 | 0.063 | 0.053 | 0.071
 | 0 | 0.062 | 0.061 | 0.081 | 0.08 | 0.108 | 0.097 | 0.097 | 0.076 | 0.094 | 0.093
 | 0 | 0.07 | 0.087 | 0.124 | 0.102 | 0.111 | 0.118 | 0.147 | 0.124 | 0.153 | 0.161
 | 0 | 0.04 | 0.059 | 0.096 | 0.094 | 0.112 | 0.131 | 0.159 | 0.148 | 0.157 | 0.165
 | 0 | -0.061 | 0.006 | 0.013 | 0.052 | 0.04 | 0.048 | 0.057 | 0.056 | 0.084 | 0.073 | 0.073
Mittelwert Druckdifferenz | 0 | 0.016333 | 0.032333 | 0.046667 | 0.056778 | 0.062 | 0.059111 | 0.080889 | 0.076333 | 0.089333 | 0.092333

Overhead shaking
Vertical stirring

## Slide 13: Remedy: Empirical compensation of wall flow: overhead shaking
Chart: Runtime [s]
| 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140
Pressure | 0 | 0.11 | 0.18 | 0.23 | 0.29 | 0.34 | 0.36 | 0.41 | 0.43 | 0.46 | 0.48
 | 0 | 0.44 | 0.59 | 0.71 | 0.76 | 0.82 | 0.87 | 0.89 | 0.94 | 0.95 | 0.98 | 1.01
 | 0 | 0.13 | 0.23 | 0.31 | 0.32 | 0.36 | 0.42 | 0.44 | 0.47 | 0.5 | 0.52
 | 0 | 0.53 | 0.71 | 0.82 | 0.89 | 0.96 | 1 | 1.06 | 1.1 | 1.12 | 1.12
 | 0 | 0.11 | 0.38 | 0.49 | 0.59 | 0.65 | 0.68 | 0.71 | 0.72 | 0.76 | 0.78
 | 0 | 0.25 | 0.41 | 0.49 | 0.57 | 0.64 | 0.65 | 0.68 | 0.68 | 0.71 | 0.73
 | 0 | 0.46 | 0.64 | 0.77 | 0.84 | 0.92 | 0.96 | 1.01 | 1.04 | 1.08 | 1.1
 | 0 | 0.28 | 0.43 | 0.48 | 0.55 | 0.57 | 0.62 | 0.66 | 0.67 | 0.69 | 0.73
 | 0 | 0.37 | 0.56 | 0.65 | 0.74 | 0.76 | 0.82 | 0.85 | 0.85 | 0.91 | 0.94
 | 0 | 0.34 | 0.54 | 0.68 | 0.75 | 0.82 | 0.89 | 0.92 | 0.94 | 0.97 | 0.99 | 1.01
Mittelwert Druckdifferenz | 0 | 0.32333 | 0.49889 | 0.6 | 0.66778 | 0.72222 | 0.76778 | 0.80222 | 0.82333 | 0.85444 | 0.87667
neue Korrektur | 0 | 0.29167 | 0.46667 | 0.58333 | 0.66667 | 0.72917 | 0.77778 | 0.81667 | 0.84848 | 0.875 | 0.89744 | 0.91667 | 0.93333 | 0.94792 | 0.96078 | 0.97222 | 0.98246 | 0.99167 | 1 | 1.0076 | 1.0145 | 1.0208 | 1.0267 | 1.0321 | 1.037 | 1.0417 | 1.046 | 1.05

## Slide 14: Remedy: Empirical compensation of wall flow: vertical stirring
Chart: Runtime [s]
| 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140
 | 0 | 0.017 | 0.022 | 0.037 | 0.033 | 0.021 | 0.017 | 0.034 | 0.052 | 0.06 | 0.047 | 0.046
 | 0 | 0.009 | 0.019 | -0.003 | 0.035 | 0.044 | 0.023 | 0.032 | 0.04 | 0.04 | 0.038
 | 0 | -0.012 | -0.003 | 0.015 | 0.032 | 0.031 | 0.028 | 0.047 | 0.027 | 0.024 | 0.042 | 0.022
 | 0 | 0.011 | 0.029 | 0.038 | 0.055 | 0.034 | 0.053 | 0.121 | 0.12 | 0.129 | 0.159
 | 0 | 0.011 | 0.011 | 0.019 | 0.028 | 0.057 | 0.017 | 0.034 | 0.044 | 0.063 | 0.053 | 0.071
 | 0 | 0.062 | 0.061 | 0.081 | 0.08 | 0.108 | 0.097 | 0.097 | 0.076 | 0.094 | 0.093
 | 0 | 0.07 | 0.087 | 0.124 | 0.102 | 0.111 | 0.118 | 0.147 | 0.124 | 0.153 | 0.161
 | 0 | 0.04 | 0.059 | 0.096 | 0.094 | 0.112 | 0.131 | 0.159 | 0.148 | 0.157 | 0.165
 | 0 | -0.061 | 0.006 | 0.013 | 0.052 | 0.04 | 0.048 | 0.057 | 0.056 | 0.084 | 0.073 | 0.073
Mittelwert Druckdifferenz | 0 | 0.016333 | 0.032333 | 0.046667 | 0.056778 | 0.062 | 0.059111 | 0.080889 | 0.076333 | 0.089333 | 0.092333

Korrektur | 0 | 0.02446 | 0.041153 | 0.053475 | 0.063051 | 0.070771 | 0.077165 | 0.082574 | 0.087228 | 0.091287 | 0.094868 | 0.098059 | 0.10092 | 0.10352 | 0.10588 | 0.10804 | 0.11003 | 0.11186 | 0.11357 | 0.11515 | 0.11664 | 0.11802 | 0.11933 | 0.12056 | 0.12172 | 0.12281 | 0.12385 | 0.12484

## Slide 15: Causes and remedies
   Mixing (overhead shaking vs. stirring)
   Temperature effects
   Errors in soil mass (air dry vs. oven-dry )
   Organic matter destruction
   Mass of dispersant

## Slide 16: Initial temperature drift in PARIO measurements
Chart: Time (s)
| 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180 | 190 | 200 | 210 | 220 | 230 | 240 | 250 | 260 | 270 | 280 | 290 | 300 | 310 | 320 | 330 | 340 | 350 | 360 | 370 | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500
Temperature [°C] | 0 | 0.019 | 0.035 | 0.049 | 0.061 | 0.071 | 0.079 | 0.086 | 0.092 | 0.096 | 0.1 | 0.103 | 0.106 | 0.108 | 0.11 | 0.111 | 0.112 | 0.113 | 0.114 | 0.114 | 0.115 | 0.116 | 0.116 | 0.117 | 0.117 | 0.117 | 0.118 | 0.118 | 0.118 | 0.118 | 0.119 | 0.119 | 0.119 | 0.119 | 0.119 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12 | 0.121 | 0.121 | 0.121 | 0.121 | 0.121 | 0.121
Temperature [°C] | 0 | 0.018 | 0.036 | 0.052 | 0.065 | 0.076 | 0.085 | 0.091 | 0.097 | 0.102 | 0.106 | 0.109 | 0.111 | 0.113 | 0.115 | 0.116 | 0.117 | 0.118 | 0.119 | 0.119 | 0.12 | 0.12 | 0.12 | 0.121 | 0.121 | 0.121 | 0.121 | 0.121 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122 | 0.122
Temperature [°C] | 0 | 0.002 | 0.006 | 0.009 | 0.012 | 0.015 | 0.017 | 0.018 | 0.02 | 0.021 | 0.022 | 0.023 | 0.024 | 0.024 | 0.025 | 0.026 | 0.026 | 0.026 | 0.027 | 0.027 | 0.027 | 0.027 | 0.028 | 0.028 | 0.028 | 0.029 | 0.029 | 0.029 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.031 | 0.031 | 0.031 | 0.032 | 0.032 | 0.032 | 0.032 | 0.033 | 0.033 | 0.033 | 0.033 | 0.033 | 0.034 | 0.034 | 0.034
Temperature [°C] | 0 | -0.065 | -0.112 | -0.151 | -0.182 | -0.207 | -0.228 | -0.244 | -0.258 | -0.269 | -0.278 | -0.284 | -0.29 | -0.294 | -0.298 | -0.3 | -0.303 | -0.304 | -0.306 | -0.307 | -0.307 | -0.308 | -0.308 | -0.308 | -0.309 | -0.308 | -0.308 | -0.308 | -0.308 | -0.307 | -0.307 | -0.307 | -0.307 | -0.306 | -0.306 | -0.305 | -0.305 | -0.305 | -0.304 | -0.303 | -0.303 | -0.303 | -0.302 | -0.302 | -0.301 | -0.301 | -0.3 | -0.3 | -0.299
Temperature [°C] | 0 | -0.03 | -0.052 | -0.071 | -0.086 | -0.098 | -0.107 | -0.115 | -0.121 | -0.126 | -0.13 | -0.132 | -0.135 | -0.137 | -0.138 | -0.139 | -0.14 | -0.14 | -0.141 | -0.141 | -0.141 | -0.141 | -0.141 | -0.141 | -0.141 | -0.141 | -0.141 | -0.141 | -0.141 | -0.14 | -0.14 | -0.14 | -0.14 | -0.139 | -0.14 | -0.139 | -0.139 | -0.138 | -0.138 | -0.138 | -0.138 | -0.137 | -0.137 | -0.137 | -0.137 | -0.136 | -0.136 | -0.136 | -0.135
Temperature [°C] | 0 | -0.071 | -0.123 | -0.167 | -0.203 | -0.232 | -0.256 | -0.276 | -0.291 | -0.303 | -0.314 | -0.321 | -0.328 | -0.333 | -0.337 | -0.34 | -0.343 | -0.345 | -0.346 | -0.347 | -0.349 | -0.349 | -0.35 | -0.35 | -0.35 | -0.35 | -0.351 | -0.351 | -0.351 | -0.351 | -0.35 | -0.35 | -0.35 | -0.35 | -0.349 | -0.349 | -0.349 | -0.349 | -0.348 | -0.348 | -0.347 | -0.347 | -0.347 | -0.346 | -0.346 | -0.345 | -0.345 | -0.344 | -0.344
Temperature [°C] | 0 | 0.006 | 0.013 | 0.019 | 0.025 | 0.03 | 0.033 | 0.037 | 0.039 | 0.041 | 0.043 | 0.045 | 0.045 | 0.047 | 0.047 | 0.048 | 0.049 | 0.049 | 0.05 | 0.05 | 0.051 | 0.051 | 0.051 | 0.052 | 0.052 | 0.052 | 0.052 | 0.053 | 0.053 | 0.053 | 0.053 | 0.053 | 0.053 | 0.053 | 0.054 | 0.054 | 0.054 | 0.054 | 0.054 | 0.054 | 0.054 | 0.054 | 0.054 | 0.055 | 0.054 | 0.055 | 0.055 | 0.055 | 0.055
Temperature [°C] | 0 | 0.004 | 0.009 | 0.013 | 0.017 | 0.02 | 0.022 | 0.024 | 0.026 | 0.027 | 0.027 | 0.028 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.029 | 0.028 | 0.028 | 0.028 | 0.028 | 0.028 | 0.028 | 0.028 | 0.027 | 0.027 | 0.027 | 0.027 | 0.027 | 0.027 | 0.026 | 0.026 | 0.026 | 0.026 | 0.026 | 0.026 | 0.026 | 0.026 | 0.026 | 0.026 | 0.026 | 0.025 | 0.025 | 0.025
Temperature [°C] | 0 | -0.008 | -0.013 | -0.016 | -0.019 | -0.021 | -0.023 | -0.024 | -0.025 | -0.026 | -0.027 | -0.027 | -0.027 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.028 | -0.027 | -0.027 | -0.027 | -0.027 | -0.027 | -0.027 | -0.026 | -0.027 | -0.027 | -0.026 | -0.027 | -0.026 | -0.026 | -0.026 | -0.026
Temperature [°C] | 0 | -0.014 | -0.024 | -0.031 | -0.036 | -0.04 | -0.043 | -0.046 | -0.048 | -0.05 | -0.051 | -0.052 | -0.053 | -0.053 | -0.054 | -0.054 | -0.054 | -0.054 | -0.054 | -0.054 | -0.054 | -0.054 | -0.054 | -0.054 | -0.054 | -0.054 | -0.054 | -0.054 | -0.053 | -0.053 | -0.053 | -0.053 | -0.053 | -0.053 | -0.053 | -0.052 | -0.052 | -0.052 | -0.052 | -0.051 | -0.051 | -0.051 | -0.051 | -0.051 | -0.05 | -0.05 | -0.05 | -0.05 | -0.049
Temperature [°C] | 0 | -0.026 | -0.043 | -0.055 | -0.066 | -0.074 | -0.082 | -0.087 | -0.092 | -0.097 | -0.1 | -0.103 | -0.105 | -0.107 | -0.109 | -0.111 | -0.112 | -0.113 | -0.114 | -0.115 | -0.116 | -0.116 | -0.117 | -0.117 | -0.117 | -0.118 | -0.118 | -0.118 | -0.119 | -0.119 | -0.119 | -0.12 | -0.119 | -0.119 | -0.12 | -0.12 | -0.12 | -0.12 | -0.12 | -0.12 | -0.12 | -0.12 | -0.12 | -0.12 | -0.12 | -0.12 | -0.119 | -0.119 | -0.119
Temperature [°C] | 0 | -0.017 | -0.029 | -0.038 | -0.045 | -0.051 | -0.056 | -0.06 | -0.063 | -0.065 | -0.067 | -0.068 | -0.069 | -0.07 | -0.071 | -0.072 | -0.072 | -0.072 | -0.073 | -0.073 | -0.073 | -0.073 | -0.073 | -0.073 | -0.073 | -0.073 | -0.073 | -0.072 | -0.072 | -0.072 | -0.072 | -0.072 | -0.071 | -0.071 | -0.071 | -0.071 | -0.071 | -0.071 | -0.07 | -0.07 | -0.07 | -0.07 | -0.069 | -0.069 | -0.069 | -0.068 | -0.068 | -0.068 | -0.067

## Slide 17: Initial temperature drift causes pressure drift
Chart: T(t) absolut
| 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180 | 190 | 200 | 210 | 220 | 230 | 240 | 250 | 260 | 270 | 280 | 290 | 300 | 310 | 320
 | 20.072 | 19.964 | 19.863 | 19.787 | 19.722 | 19.664 | 19.611 | 19.562 | 19.517 | 19.478 | 19.443 | 19.411 | 19.382 | 19.357 | 19.333 | 19.312 | 19.292 | 19.274 | 19.258 | 19.243 | 19.23 | 19.217 | 19.206 | 19.196 | 19.185 | 19.177 | 19.169 | 19.162 | 19.155 | 19.149 | 19.143 | 19.138
 | 19.967 | 19.998 | 20.031 | 20.063 | 20.093 | 20.117 | 20.138 | 20.156 | 20.172 | 20.186 | 20.2 | 20.211 | 20.222 | 20.231 | 20.24 | 20.248 | 20.255 | 20.262 | 20.268 | 20.273 | 20.278 | 20.283 | 20.287 | 20.291 | 20.295 | 20.299 | 20.302 | 20.305 | 20.308 | 20.311 | 20.314 | 20.317
 | 20.059 | 19.964 | 19.87 | 19.797 | 19.737 | 19.684 | 19.636 | 19.591 | 19.549 | 19.51 | 19.475 | 19.444 | 19.415 | 19.389 | 19.365 | 19.343 | 19.324 | 19.305 | 19.289 | 19.273 | 19.26 | 19.247 | 19.235 | 19.225 | 19.215 | 19.206 | 19.198 | 19.191 | 19.184 | 19.178 | 19.172 | 19.167
 | 20.012 | 20.046 | 20.084 | 20.113 | 20.139 | 20.161 | 20.181 | 20.198 | 20.214 | 20.228 | 20.24 | 20.251 | 20.26 | 20.269 | 20.277 | 20.284 | 20.291 | 20.297 | 20.303 | 20.309 | 20.314 | 20.319 | 20.323 | 20.327 | 20.331 | 20.335 | 20.339 | 20.342 | 20.345 | 20.348 | 20.351 | 20.354
 | 20.03 | 19.923 | 19.837 | 19.767 | 19.708 | 19.656 | 19.611 | 19.571 | 19.535 | 19.503 | 19.473 | 19.447 | 19.422 | 19.399 | 19.379 | 19.36 | 19.342 | 19.327 | 19.312 | 19.299 | 19.287 | 19.275 | 19.265 | 19.256 | 19.247 | 19.24 | 19.233 | 19.226 | 19.22 | 19.22
 | 20.128 | 20.163 | 20.194 | 20.218 | 20.238 | 20.253 | 20.267 | 20.28 | 20.291 | 20.301 | 20.31 | 20.319 | 20.327 | 20.334 | 20.34 | 20.346 | 20.351 | 20.356 | 20.361 | 20.365 | 20.369 | 20.373 | 20.376 | 20.38 | 20.383 | 20.386 | 20.388 | 20.39 | 20.393 | 20.393
 | 20.045 | 19.938 | 19.858 | 19.796 | 19.742 | 19.692 | 19.649 | 19.61 | 19.574 | 19.541 | 19.511 | 19.483 | 19.459 | 19.436 | 19.415 | 19.397 | 19.38 | 19.364 | 19.35 | 19.336 | 19.324 | 19.313 | 19.303 | 19.294 | 19.286 | 19.277 | 19.271 | 19.264 | 19.258 | 19.252 | 19.252
 | 19.943 | 19.988 | 20.031 | 20.067 | 20.099 | 20.125 | 20.147 | 20.167 | 20.183 | 20.198 | 20.211 | 20.222 | 20.232 | 20.241 | 20.25 | 20.258 | 20.265 | 20.272 | 20.278 | 20.283 | 20.289 | 20.294 | 20.3 | 20.304 | 20.308 | 20.313 | 20.317 | 20.321 | 20.325 | 20.328 | 20.331
Chart: p(t) absolut
| 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180 | 190 | 200 | 210 | 220 | 230 | 240 | 250 | 260 | 270 | 280 | 290 | 300 | 310 | 320
 | 1782.35 | 1785.2 | 1786.19 | 1786.82 | 1787.11 | 1787.22 | 1787.24 | 1787.29 | 1787.33 | 1787.41 | 1787.43 | 1787.47 | 1787.51 | 1787.6 | 1787.6 | 1787.65 | 1787.65 | 1787.7 | 1787.71 | 1787.75 | 1787.76 | 1787.77 | 1787.8 | 1787.81 | 1787.82 | 1787.82 | 1787.84 | 1787.84 | 1787.88 | 1787.86 | 1787.88 | 1787.88
 | 1779.66 | 1776.32 | 1775.42 | 1775.12 | 1774.92 | 1774.83 | 1774.74 | 1774.6 | 1774.5 | 1774.39 | 1774.29 | 1774.19 | 1774.07 | 1774.04 | 1774 | 1773.96 | 1773.93 | 1773.86 | 1773.85 | 1773.84 | 1773.81 | 1773.78 | 1773.76 | 1773.73 | 1773.73 | 1773.7 | 1773.72 | 1773.71 | 1773.7 | 1773.7 | 1773.65 | 1773.65
 | 1762.71 | 1764.37 | 1764.81 | 1765.09 | 1765.25 | 1765.3 | 1765.36 | 1765.42 | 1765.44 | 1765.46 | 1765.54 | 1765.54 | 1765.57 | 1765.58 | 1765.58 | 1765.57 | 1765.58 | 1765.57 | 1765.59 | 1765.62 | 1765.64 | 1765.64 | 1765.63 | 1765.65 | 1765.64 | 1765.65 | 1765.64 | 1765.66 | 1765.64 | 1765.65 | 1765.68 | 1765.6
 | 1760.55 | 1757.37 | 1756.28 | 1755.85 | 1755.67 | 1755.61 | 1755.54 | 1755.46 | 1755.35 | 1755.24 | 1755.14 | 1755.06 | 1755 | 1754.89 | 1754.9 | 1754.86 | 1754.8 | 1754.79 | 1754.76 | 1754.75 | 1754.73 | 1754.7 | 1754.75 | 1754.63 | 1754.66 | 1754.62 | 1754.71 | 1754.62 | 1754.52 | 1754.57 | 1754.57 | 1754.56
 | 1742.17 | 1743.57 | 1744.07 | 1744.41 | 1744.57 | 1744.69 | 1744.75 | 1744.81 | 1744.81 | 1744.85 | 1744.88 | 1744.91 | 1744.91 | 1744.92 | 1744.95 | 1744.95 | 1744.96 | 1744.97 | 1744.98 | 1744.98 | 1745 | 1745 | 1745 | 1745.02 | 1745.02 | 1745 | 1745.03 | 1745.04 | 1745.04 | 1745.05 | 1745.13 | 1745.02
 | 1735.76 | 1734.08 | 1733.48 | 1733.12 | 1732.91 | 1732.75 | 1732.65 | 1732.57 | 1732.52 | 1732.46 | 1732.38 | 1732.32 | 1732.28 | 1732.24 | 1732.22 | 1732.18 | 1732.12 | 1732.1 | 1732.09 | 1732.06 | 1732.04 | 1732.02 | 1732.01 | 1732.01 | 1732.02 | 1731.99 | 1731.97 | 1731.98 | 1731.97 | 1731.97
 | 1723.52 | 1724.17 | 1724.45 | 1724.58 | 1724.65 | 1724.73 | 1724.71 | 1724.78 | 1724.8 | 1724.83 | 1724.84 | 1724.87 | 1724.88 | 1724.89 | 1724.92 | 1724.91 | 1724.91 | 1724.93 | 1724.92 | 1724.96 | 1724.96 | 1724.98 | 1724.97 | 1724.98 | 1724.97 | 1724.98 | 1724.98 | 1724.98 | 1725 | 1725
 | 1716.85 | 1712.65 | 1711.78 | 1711.38 | 1711.16 | 1711.01 | 1710.92 | 1710.87 | 1710.82 | 1710.76 | 1710.68 | 1710.64 | 1710.59 | 1710.53 | 1710.5 | 1710.46 | 1710.44 | 1710.41 | 1710.38 | 1710.34 | 1710.32 | 1710.32 | 1710.29 | 1710.29 | 1710.28 | 1710.25 | 1710.24 | 1710.24 | 1710.25 | 1710.24 | 1710.24
 | 1696.28 | 1696.94 | 1697.33 | 1697.59 | 1697.72 | 1697.81 | 1697.85 | 1697.9 | 1697.93 | 1697.98 | 1698 | 1698.02 | 1698.04 | 1698.06 | 1698.1 | 1698.13 | 1698.12 | 1698.12 | 1698.13 | 1698.17 | 1698.17 | 1698.18 | 1698.17 | 1698.2 | 1698.2 | 1698.21 | 1698.21 | 1698.21 | 1698.23 | 1698.22 | 1698.22

## Slide 18: Forced temperature drift causes pressure drift
Chart: pressure deviation
| 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180 | 190 | 200 | 210 | 220 | 230 | 240 | 250 | 260 | 270 | 280 | 290 | 300
 | -2.66 | -1.67 | -1.04 | -0.75 | -0.64 | -0.62 | -0.57 | -0.53 | -0.45 | -0.43 | -0.39 | -0.35 | -0.26 | -0.26 | -0.21 | -0.21 | -0.16 | -0.15 | -0.11 | -0.1 | -0.09 | -0.06 | -0.05 | -0.04 | -0.04 | -0.02 | -0.02 | 0.02 | 0
 | 2.62 | 1.72 | 1.42 | 1.22 | 1.13 | 1.04 | 0.9 | 0.8 | 0.69 | 0.59 | 0.49 | 0.37 | 0.34 | 0.3 | 0.26 | 0.23 | 0.16 | 0.15 | 0.14 | 0.11 | 0.08 | 0.06 | 0.03 | 0.03 | 0 | 0.02 | 0.01 | 0 | 0
 | -1.28 | -0.84 | -0.56 | -0.4 | -0.35 | -0.29 | -0.23 | -0.21 | -0.19 | -0.11 | -0.11 | -0.08 | -0.07 | -0.07 | -0.08 | -0.07 | -0.08 | -0.06 | -0.03 | -0.01 | -0.01 | -0.02 | 0 | -0.01 | 0 | -0.01 | 0.01 | -0.01 | 0
 | 2.8 | 1.71 | 1.28 | 1.1 | 1.04 | 0.97 | 0.89 | 0.78 | 0.67 | 0.57 | 0.49 | 0.43 | 0.32 | 0.33 | 0.29 | 0.23 | 0.22 | 0.19 | 0.18 | 0.16 | 0.13 | 0.18 | 0.06 | 0.09 | 0.05 | 0.14 | 0.05 | -0.05 | 0
 | -1.48 | -0.98 | -0.64 | -0.48 | -0.36 | -0.3 | -0.24 | -0.24 | -0.2 | -0.17 | -0.14 | -0.14 | -0.13 | -0.1 | -0.1 | -0.09 | -0.08 | -0.07 | -0.07 | -0.05 | -0.05 | -0.05 | -0.03 | -0.03 | -0.05 | -0.02 | -0.01 | -0.01 | 0
 | 3.79 | 2.11 | 1.51 | 1.15 | 0.94 | 0.78 | 0.68 | 0.6 | 0.55 | 0.49 | 0.41 | 0.35 | 0.31 | 0.27 | 0.25 | 0.21 | 0.15 | 0.13 | 0.12 | 0.09 | 0.07 | 0.05 | 0.04 | 0.04 | 0.05 | 0.02 | 0 | 0.01 | 0
 | -1.48 | -0.83 | -0.55 | -0.42 | -0.35 | -0.27 | -0.29 | -0.22 | -0.2 | -0.17 | -0.16 | -0.13 | -0.12 | -0.11 | -0.08 | -0.09 | -0.09 | -0.07 | -0.08 | -0.04 | -0.04 | -0.02 | -0.03 | -0.02 | -0.03 | -0.02 | -0.02 | -0.02 | 0
 | 2.41 | 1.54 | 1.14 | 0.92 | 0.77 | 0.68 | 0.63 | 0.58 | 0.52 | 0.44 | 0.4 | 0.35 | 0.29 | 0.26 | 0.22 | 0.2 | 0.17 | 0.14 | 0.1 | 0.08 | 0.08 | 0.05 | 0.05 | 0.04 | 0.01 | 0 | 0 | 0.01 | 0
 | -1.95 | -1.29 | -0.9 | -0.64 | -0.51 | -0.42 | -0.38 | -0.33 | -0.3 | -0.25 | -0.23 | -0.21 | -0.19 | -0.17 | -0.13 | -0.1 | -0.11 | -0.11 | -0.1 | -0.06 | -0.06 | -0.05 | -0.06 | -0.03 | -0.03 | -0.02 | -0.02 | -0.02 | 0
Chart: temperature deviation
| 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180 | 190 | 200 | 210 | 220 | 230 | 240 | 250 | 260 | 270 | 280 | 290 | 300
 | -0.57 | -0.512 | -0.453 | -0.4 | -0.355 | -0.316 | -0.283 | -0.253 | -0.227 | -0.204 | -0.184 | -0.167 | -0.152 | -0.137 | -0.125 | -0.113 | -0.101 | -0.09 | -0.078 | -0.068 | -0.058 | -0.048 | -0.04 | -0.032 | -0.025 | -0.018 | -0.012 | -0.006 | 0
 | 0.815 | 0.714 | 0.638 | 0.573 | 0.515 | 0.462 | 0.413 | 0.368 | 0.329 | 0.294 | 0.262 | 0.233 | 0.208 | 0.184 | 0.163 | 0.143 | 0.125 | 0.109 | 0.094 | 0.081 | 0.068 | 0.057 | 0.047 | 0.036 | 0.028 | 0.02 | 0.013 | 0.006 | 0
 | -0.313 | -0.28 | -0.248 | -0.218 | -0.194 | -0.173 | -0.155 | -0.139 | -0.125 | -0.111 | -0.1 | -0.089 | -0.08 | -0.071 | -0.063 | -0.056 | -0.049 | -0.043 | -0.038 | -0.033 | -0.028 | -0.024 | -0.02 | -0.016 | -0.012 | -0.009 | -0.006 | -0.003 | 0
 | 0.786 | 0.692 | 0.619 | 0.559 | 0.506 | 0.458 | 0.413 | 0.371 | 0.332 | 0.297 | 0.266 | 0.237 | 0.211 | 0.187 | 0.165 | 0.146 | 0.127 | 0.111 | 0.095 | 0.082 | 0.069 | 0.057 | 0.047 | 0.037 | 0.028 | 0.02 | 0.013 | 0.006 | 0
 | -0.302 | -0.264 | -0.235 | -0.209 | -0.187 | -0.167 | -0.15 | -0.134 | -0.12 | -0.108 | -0.097 | -0.088 | -0.079 | -0.071 | -0.064 | -0.057 | -0.051 | -0.045 | -0.039 | -0.034 | -0.029 | -0.025 | -0.021 | -0.017 | -0.013 | -0.009 | -0.006 | -0.003 | 0
 | 0.81 | 0.703 | 0.617 | 0.547 | 0.488 | 0.436 | 0.391 | 0.351 | 0.315 | 0.283 | 0.253 | 0.227 | 0.202 | 0.179 | 0.159 | 0.14 | 0.122 | 0.107 | 0.092 | 0.079 | 0.067 | 0.055 | 0.045 | 0.036 | 0.027 | 0.02 | 0.013 | 0.006 | 0
 | -0.265 | -0.23 | -0.199 | -0.175 | -0.155 | -0.14 | -0.126 | -0.113 | -0.102 | -0.092 | -0.083 | -0.074 | -0.066 | -0.059 | -0.053 | -0.047 | -0.042 | -0.037 | -0.032 | -0.028 | -0.024 | -0.02 | -0.017 | -0.013 | -0.01 | -0.007 | -0.005 | -0.003 | 0
 | 0.686 | 0.606 | 0.544 | 0.49 | 0.44 | 0.397 | 0.358 | 0.322 | 0.289 | 0.259 | 0.231 | 0.207 | 0.184 | 0.163 | 0.145 | 0.128 | 0.112 | 0.098 | 0.084 | 0.072 | 0.061 | 0.051 | 0.042 | 0.034 | 0.025 | 0.019 | 0.012 | 0.006 | 0
 | -0.382 | -0.337 | -0.294 | -0.258 | -0.226 | -0.2 | -0.178 | -0.158 | -0.142 | -0.127 | -0.114 | -0.103 | -0.093 | -0.084 | -0.075 | -0.067 | -0.06 | -0.053 | -0.047 | -0.042 | -0.036 | -0.031 | -0.025 | -0.021 | -0.017 | -0.012 | -0.008 | -0.004 | 0

## Slide 19: Pressure deviation vs. temperature deviation
Chart: Dp
| 0.81499999999999773 | 0.71399999999999864 | 0.63799999999999812 | 0.5730000000000004 | 0.51500000000000057 | 0.46199999999999974 | 0.41300000000000026 | 0.36799999999999855 | 0.32900000000000063 | 0.29400000000000048 | 0.26200000000000045 | 0.23300000000000054 | 0.20799999999999841 | 0.1839999999999975 | 0.16300000000000026 | 0.14300000000000068 | 0.125 | 0.10899999999999821 | 9.3999999999997641E-2 | 8.0999999999999517E-2 | 6.799999999999784E-2 | 5.6999999999998607E-2 | 4.7000000000000597E-2 | 3.5999999999997812E-2 | 2.7999999999998693E-2 | 1.9999999999999574E-2 | 1.2999999999998124E-2 | 6.0000000000002274E-3 | 0
 | 2.62 | 1.72 | 1.42 | 1.22 | 1.13 | 1.04 | 0.9 | 0.8 | 0.69 | 0.59 | 0.49 | 0.37 | 0.34 | 0.3 | 0.26 | 0.23 | 0.16 | 0.15 | 0.14 | 0.11 | 0.08 | 0.06 | 0.03 | 0.03 | 0 | 0.02 | 0.01 | 0 | 0
 | -1.28 | -0.84 | -0.56 | -0.4 | -0.35 | -0.29 | -0.23 | -0.21 | -0.19 | -0.11 | -0.11 | -0.08 | -0.07 | -0.07 | -0.08 | -0.07 | -0.08 | -0.06 | -0.03 | -0.01 | -0.01 | -0.02 | 0 | -0.01 | 0 | -0.01 | 0.01 | -0.01 | 0
 | 2.8 | 1.71 | 1.28 | 1.1 | 1.04 | 0.97 | 0.89 | 0.78 | 0.67 | 0.57 | 0.49 | 0.43 | 0.32 | 0.33 | 0.29 | 0.23 | 0.22 | 0.19 | 0.18 | 0.16 | 0.13 | 0.18 | 0.06 | 0.09 | 0.05 | 0.14 | 0.05 | -0.05 | 0
 | -1.48 | -0.98 | -0.64 | -0.48 | -0.36 | -0.3 | -0.24 | -0.24 | -0.2 | -0.17 | -0.14 | -0.14 | -0.13 | -0.1 | -0.1 | -0.09 | -0.08 | -0.07 | -0.07 | -0.05 | -0.05 | -0.05 | -0.03 | -0.03 | -0.05 | -0.02 | -0.01 | -0.01 | 0
 | 3.79 | 2.11 | 1.51 | 1.15 | 0.94 | 0.78 | 0.68 | 0.6 | 0.55 | 0.49 | 0.41 | 0.35 | 0.31 | 0.27 | 0.25 | 0.21 | 0.15 | 0.13 | 0.12 | 0.09 | 0.07 | 0.05 | 0.04 | 0.04 | 0.05 | 0.02 | 0 | 0.01 | 0
 | -1.48 | -0.83 | -0.55 | -0.42 | -0.35 | -0.27 | -0.29 | -0.22 | -0.2 | -0.17 | -0.16 | -0.13 | -0.12 | -0.11 | -0.08 | -0.09 | -0.09 | -0.07 | -0.08 | -0.04 | -0.04 | -0.02 | -0.03 | -0.02 | -0.03 | -0.02 | -0.02 | -0.02 | 0
 | 2.41 | 1.54 | 1.14 | 0.92 | 0.77 | 0.68 | 0.63 | 0.58 | 0.52 | 0.44 | 0.4 | 0.35 | 0.29 | 0.26 | 0.22 | 0.2 | 0.17 | 0.14 | 0.1 | 0.08 | 0.08 | 0.05 | 0.05 | 0.04 | 0.01 | 0 | 0 | 0.01 | 0
 | -1.95 | -1.29 | -0.9 | -0.64 | -0.51 | -0.42 | -0.38 | -0.33 | -0.3 | -0.25 | -0.23 | -0.21 | -0.19 | -0.17 | -0.13 | -0.1 | -0.11 | -0.11 | -0.1 | -0.06 | -0.06 | -0.05 | -0.06 | -0.03 | -0.03 | -0.02 | -0.02 | -0.02 | 0
DT

## Slide 20: Empirical correction for temperature drift
Chart: (range -0.4 to + 0.8)
DT
Dp

## Slide 21: Remedy: Effect of the T-drift correction
Chart: rel. pressure (t)
| 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180 | 190 | 200 | 210 | 220 | 230 | 240 | 250 | 260 | 270 | 280 | 290 | 300
 | -2.66 | -1.67 | -1.04 | -0.75 | -0.64 | -0.62 | -0.57 | -0.53 | -0.45 | -0.43 | -0.39 | -0.35 | -0.26 | -0.26 | -0.21 | -0.21 | -0.16 | -0.15 | -0.11 | -0.1 | -0.09 | -0.06 | -0.05 | -0.04 | -0.04 | -0.02 | -0.02 | 0.02 | 0
 | 2.62 | 1.72 | 1.42 | 1.22 | 1.13 | 1.04 | 0.9 | 0.8 | 0.69 | 0.59 | 0.49 | 0.37 | 0.34 | 0.3 | 0.26 | 0.23 | 0.16 | 0.15 | 0.14 | 0.11 | 0.08 | 0.06 | 0.03 | 0.03 | 0 | 0.02 | 0.01 | 0 | 0
 | -1.28 | -0.84 | -0.56 | -0.4 | -0.35 | -0.29 | -0.23 | -0.21 | -0.19 | -0.11 | -0.11 | -0.08 | -0.07 | -0.07 | -0.08 | -0.07 | -0.08 | -0.06 | -0.03 | -0.01 | -0.01 | -0.02 | 0 | -0.01 | 0 | -0.01 | 0.01 | -0.01 | 0
 | 2.8 | 1.71 | 1.28 | 1.1 | 1.04 | 0.97 | 0.89 | 0.78 | 0.67 | 0.57 | 0.49 | 0.43 | 0.32 | 0.33 | 0.29 | 0.23 | 0.22 | 0.19 | 0.18 | 0.16 | 0.13 | 0.18 | 0.06 | 0.09 | 0.05 | 0.14 | 0.05 | -0.05 | 0
 | -1.48 | -0.98 | -0.64 | -0.48 | -0.36 | -0.3 | -0.24 | -0.24 | -0.2 | -0.17 | -0.14 | -0.14 | -0.13 | -0.1 | -0.1 | -0.09 | -0.08 | -0.07 | -0.07 | -0.05 | -0.05 | -0.05 | -0.03 | -0.03 | -0.05 | -0.02 | -0.01 | -0.01 | 0
 | 3.79 | 2.11 | 1.51 | 1.15 | 0.94 | 0.78 | 0.68 | 0.6 | 0.55 | 0.49 | 0.41 | 0.35 | 0.31 | 0.27 | 0.25 | 0.21 | 0.15 | 0.13 | 0.12 | 0.09 | 0.07 | 0.05 | 0.04 | 0.04 | 0.05 | 0.02 | 0 | 0.01 | 0
 | -1.48 | -0.83 | -0.55 | -0.42 | -0.35 | -0.27 | -0.29 | -0.22 | -0.2 | -0.17 | -0.16 | -0.13 | -0.12 | -0.11 | -0.08 | -0.09 | -0.09 | -0.07 | -0.08 | -0.04 | -0.04 | -0.02 | -0.03 | -0.02 | -0.03 | -0.02 | -0.02 | -0.02 | 0
 | 2.41 | 1.54 | 1.14 | 0.92 | 0.77 | 0.68 | 0.63 | 0.58 | 0.52 | 0.44 | 0.4 | 0.35 | 0.29 | 0.26 | 0.22 | 0.2 | 0.17 | 0.14 | 0.1 | 0.08 | 0.08 | 0.05 | 0.05 | 0.04 | 0.01 | 0 | 0 | 0.01 | 0
 | -1.95 | -1.29 | -0.9 | -0.64 | -0.51 | -0.42 | -0.38 | -0.33 | -0.3 | -0.25 | -0.23 | -0.21 | -0.19 | -0.17 | -0.13 | -0.1 | -0.11 | -0.11 | -0.1 | -0.06 | -0.06 | -0.05 | -0.06 | -0.03 | -0.03 | -0.02 | -0.02 | -0.02 | 0
Chart: rel. pressure (t < 30 s)
| 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180 | 190 | 200 | 210 | 220 | 230 | 240 | 250 | 260 | 270 | 280 | 290 | 300
 | -0.32292 | -0.12074 | -0.043097 | 0.05187 | 0.10519 | 0.087917 | 0.09714 | 0.080209 | 0.062202 | 0.035228 | -0.020872 | 0.0019761 | 0.010392 | 0.010671 | 0.017097 | -0.021781 | -0.0054353 | 0.0081737 | -0.0021703 | -0.013199 | -0.017624 | -0.033791 | -0.018868 | -0.038164 | -0.0075441 | -0.0082911 | -0.0090443 | -0.0011
 | 0.25264 | 0.24455 | 0.19713 | 0.11656 | 0.083405 | 0.076628 | 0.045796 | 0.026918 | 0.072497 | 0.048135 | 0.055801 | 0.048839 | 0.032901 | 0.0094813 | 0.0082431 | -0.012585 | -0.001581 | 0.021098 | 0.033919 | 0.026863 | 0.011296 | 0.02579 | 0.010337 | 0.014928 | 0.00089652 | 0.016883 | -0.0071146 | -0.0011
 | -0.16078 | -0.16887 | -0.11449 | -0.012562 | 0.045454 | 0.077917 | 0.069922 | 0.053104 | 0.035248 | 0.026228 | 0.030467 | -0.024242 | 0.034474 | 0.036923 | 0.011757 | 0.034839 | 0.031338 | 0.046631 | 0.046343 | 0.035364 | 0.10238 | -0.0037908 | 0.039786 | 0.011836 | 0.11246 | 0.031709 | -0.059044 | -0.0011
 | -0.04097 | 0.068 | 0.064845 | 0.073503 | 0.049914 | 0.049945 | 0.0013548 | 0.0041497 | 0.0056342 | 0.011855 | -0.0061383 | -0.01298 | 0.0029006 | -0.0088765 | -0.010178 | -0.0095299 | -0.0086094 | -0.017449 | -0.0046556 | -0.011735 | -0.017318 | -0.0028384 | -0.0083045 | -0.033724 | -0.0091035 | -0.0031167 | -0.0071146 | -0.0011
 | 0.15634 | 0.070217 | -0.025083 | -0.063128 | -0.088983 | -0.078412 | -0.062112 | -0.026783 | -0.012447 | -0.024677 | -0.027983 | -0.015695 | -0.0098347 | 0.0081104 | 0.0023941 | -0.026748 | -0.022228 | -0.0087551 | -0.019209 | -0.021766 | -0.024834 | -0.021057 | -0.0088681 | 0.013168 | -0.0075441 | -0.018291 | 0.00095573 | -0.0011
 | -0.15634 | -0.058516 | -0.038473 | -0.043372 | -0.011237 | -0.070461 | -0.032833 | -0.037596 | -0.028292 | -0.035627 | -0.021892 | -0.025563 | -0.026994 | -0.0064445 | -0.025611 | -0.033057 | -0.020348 | -0.037502 | -0.0031368 | -0.0087035 | 0.0057903 | -0.0083045 | -0.0037237 | -0.017762 | -0.011781 | -0.014451 | -0.017115 | -0.0011
 | 0.14837 | -0.025522 | -0.088522 | -0.10899 | -0.092992 | -0.048859 | -0.013256 | 0.0037544 | -0.0080496 | 0.013438 | 0.014041 | 0.00039193 | 0.010671 | 0.0035419 | 0.013142 | 0.0097179 | 0.001978 | -0.016643 | -0.018968 | -0.003241 | -0.01929 | -0.006974 | -0.006182 | -0.024173 | -0.026221 | -0.01697 | 0.00095573 | -0.0011
 | 0.53895 | 0.34443 | 0.24647 | 0.13725 | 0.076606 | 0.013992 | -0.012991 | -0.035224 | -0.027817 | -0.040468 | -0.045435 | -0.046294 | -0.043757 | -0.020133 | -0.0038909 | -0.025389 | -0.036444 | -0.035611 | -0.0030573 | -0.011789 | -0.0089173 | -0.027318 | -0.0028384 | -0.0083045 | -0.0050721 | -0.010443 | -0.015783 | -0.0011

## Slide 22: Causes and remedies
   Mixing (overhead shaking vs. stirring)
   Temperature effects
   Errors in soil mass (air dry vs. oven-dry )
   Organic matter destruction
   Mass of dispersant

## Slide 23: Soil mass must be dry net mass of sedimenting particles
When calculating the mineral particle weight, pay attention to the water content of the air-dry mass and, if necessary, the content of organic matter to be destroyed.
)
)
Example calculation: ;
Mass | 25 g | 23.8 g | 23.2 g
% silt | 36.4 % | 33.7 5 | 32.2 %
% clay | 55.0 % | 58.0 % | 59.5 %
Mass | 25 g | 23.8 g | 23.2 g
% silt | 36.4 % | 33.7 5 | 32.2 %
% clay | 55.0 % | 58.0 % | 59.5 %
Notes: PARIO derives the silt content from the pressure changes as absolute mass and relates this to the original mass. Thus, any error in the dry soil mass strongly influences the analysis. An incorrect total mass can be caused, for example, by using air-dry material and entering this in the PARIO Control interface as total (oven) dry mass. To avoid this error, the gravimetric water content in the air-dry mass must be known to recalculate the oven-dry mass. Another source of error can be caused by not considering the mass loss by destruction of organic matter during the preparation of the suspension. If organic matter is destroyed, it also must be subtracted from total air-dry mass

## Slide 24: Causes and remedies
   Mixing (overhead shaking vs. stirring)
   Temperature effects
   Errors in soil mass (air dry vs. oven-dry )
   Organic matter destruction
   Mass of dispersant

## Slide 25: Problem: Gas formation by incomplete oxidation
Chart: Runtime [s]

## Slide 26: Problem: Gas formation by incomplete oxidation
Chart: Runtime [s]

## Slide 27: Problem: Gas formation by incomplete oxidation
Chart: Runtime [s]

## Slide 28: Remedy
If H2O2 is used : Ensure complete oxidation by adding enough heat and providing enough time.
Problem is particularly virulent for clay-humus complexes.

## Slide 29: Causes and remedies
   Mixing (overhead shaking vs. stirring)
   Temperature effects
   Errors in soil mass (air dry vs. oven-dry )
   Organic matter destruction
   Mass of dispersant

## Slide 30: Mass of dispersant: problem and solution
Example Na2P4O7  10 H2O (tetra-sodium diphosphate decahydrate, )
Original concentration of dispersant salt in dispersant solution | 44.6 g/L
dry concentration of dispersion salt in dispersant solution | 26.6 g/L
Errors in the assumed salt concentration is of less relevance in mode PARIO classic, but of utmost importance in PARIOplus.

## Slide 31: Overview
Intro
Accuracy of PARIO
ISP+
Test with real data

## Slide 32: ISP+ ?
    ISP+ (read “ISP plus”) is an extension of the experimental protocol.
    At some time in a ISP+ measurement, part of the suspension is released laterally from the sedimentation, collected in a beaker, and oven-dried.
    The dry mass of the collected soil particles is included into the objective function of the inverse problem.
ISP measurement
Release of suspension

## Slide 33: Object function in ISP+
    Integration of the collected dry mass, , into the objective function of the inverse problem regularizes the inverse problem of estimating the parameter vector
    It reduces the uncertainty of the identified PSD in the fine particle range (clay and fine silt fractions).

## Slide 34: Overview
Intro
Theory ISP
Theory ISP+
Test with real data

## Slide 35: PARIOplus Test Soil 1

## Slide 36: PARIOplus Test Soil 2

## Slide 37: PARIOplus Test Soil 3

## Slide 38: PARIOplus Test Soil 4

## Slide 39: PARIOplus Test Soil 5

## Slide 40: Results of replicate analyses for five test soils.

## Slide 41: Further problems: Foam

## Slide 42: Foam and/or gas bubble formation
Chart: Runtime [s]

## Slide 43: Conclusions
Temperature effects, water flow from cylinder and PARIO walls cause uncertainties in the early pressure recordings of PARIO, leading to uncertain determination of the coarse silt fraction in ISP.
Errors in coarse silt fractions pause through to complementary errors in the clay fraction.
The ISP+ method solves the problem of uncertainty of the clay fraction determination.
Also, measurement time for the particle size distribution analysis can be shortened down to ~2 hrs.
ISP measurement
Release of suspension

## Slide 44
Supplementary material
Theory of the ISP method

## Slide 45: Theory (1): Settling of particles by Stokes‘ law
= 9.81 m s-2 is the acceleration due to gravity rp [kg m-3] is the density of the suspended particles rsol [kg m-3] is the density of the liquid
 [kg s-1 m-1] is the dynamic viscosity of the liquid

## Slide 46: Theory (2): Diameter of particles reaching z after settling time t
 | (1)
 | (1)
The equivalent diameter of the particles which have reached depth z after time t is calculated by noting that z=v t

## Slide 47: Theory (3): relating mass fraction F to hydrodynamic diameter D
 | (2)
 | (2)
The corresponding mass fraction is given by the cumulative particle mass distribution function F(D) [-] which is defined as:
is a dummy variable for the particle diameter is the mass of particles with an effective diameter smaller or equal to D is the statistical density function describing the particle-size distribution.

## Slide 48: Theory (4): Pressure in a suspension at a detain depth z
[m] is the vertical coordinate, positive downward
[kg m-3] is suspension density above the measurement depth
= 9.81 m s-2 is the acceleration due to gravity

## Slide 49: Theory (5): Initial density of suspension
[kg m-3] is the suspension density right after mixing
[kg m-3] is the liquid density
[kg m-3] is the particle density
[kg m-3] is the additional density caused by suspended particles
[m3] is the volume of particles
[m3] is the total volume of the suspension [kg] is the mass of particles

## Slide 50: Theory (6): relating density at z to F(D)
 | (3)
 | (3)
The density of the suspension at depth z and at time t is obtained by extending Eq. (2) by the mass fraction given by F(D)

## Slide 51: Theory (7): relating pressure at z to F(D)
 | (4)
 | (4)
The integral in Eq. (4) can be solved analytically for some distribution functions. In the general case, i.e. for arbitrary distribution functions, the calculation of the pressure with Eq. (4) must be done by numerical integration.
This relates the integral of the cumulative mass distribution function to the measured pressure at depth .

## Slide 52: Estimation of F(D) by inverse modelling
is parameterized as Hermitian spline with 12 support points, ranging from at 0.01 µm to at 2000 µm.
Object function contains measured pressure data and measured data for sand

## Slide 53: Capillary, film and vapour flow during transient evaporation: Inverse simulation of lab experiments using pressure head and humidity data
Introduction on Evaporation
Improved simple parametrizations of soil hydraulic properties (saturation to complete dryness)
Soil hydraulic properties identified by numerical inversion of evaporation experiments

## Slide 54: Soil Hydraulic Properties: An example from Mualem‘s soil catalogue (1976)
capillary
film & corner
vapor
Peters, A., Water Resources Research, 2013
Iden S.C. and W. Durner, Water Resources Research, 2014
Peters, A., Water Resources Research, 2014
Notes: Soil hydraulic properties are required to predict evaporation from soil
In particular during stage-two evaporation, soil water retention and hydraulic conductivity curve are needed
They are also needed for predicting the duration of stage-one (atmosphere-driven) evaporation
For K(h) three regions can be distinguished: capillary. film and corner, vapor flow

## Slide 55: Parametrization of soil hydraulic properties: the PDI model (Peters, Durner, Iden)
7 parameters
only one additional parameter compared to the van Genuchten model ()
closed-form equation for
slope of relative conductivity function for film flow can be set to constant value under some assumptions (Tokunaga, 2009; Peters, 2013)

## Slide 56: Inverse modeling of evaporation experiments: Goodness-of-Fit and colinearity (Belsley, 1991)
Model | Number of parameters | Belsley-Index
VG | 5 | 42
VG-VAP | 6 | 11
PDI-NOFILM | 6 | 28
PDI | 7 | 20
PDI-A | 8 | 26
Note: The number of estimated is NOT a good proxy of parameter identifiability!

## Slide 57: Overview
Intro
Theory ISP
Theory ISP+
Test with real data

## Slide 58: In a nutshell: The ISP theory
Here, is the suspension pressure, measured at time in depth , is the liquid density, is gravitational acceleration, is the density increase by the presence of all particles in the suspension, and is the cumulative particle mass distribution function of particles up to diameter . The term expresses the constant pressure component provided by the liquid, the most right term the time-variable part.
The integral can be solved analytically for some distribution functions. In the general case, i.e. for arbitrary distribution functions, the calculation of the pressure must be done by numerical integration.
In the ISP method, the integral of the cumulative mass distribution function is related to the measured pressure at depth by

## Slide 59: ISP: Estimation of F(D) by inverse modelling
is represented by an Hermitian Spline that covers the range at 0.01 µm to at 2000 µm with 12 support points.
The spline‘s parameters are obtained by matching a simulated pressure decrease with the observed one, i.e., by minimizing an objective function that contains the measured pressure data and the sand fractions data.
Details are given in Durner, W., S.C. Iden, and G. von Unold (2017): The integral suspension pressure method (ISP) for precise particle-size analysis by gravitational sedimentation, Water Resources Research, 53, 33-48, doi:10.1002/2016WR019830 (open access).

## Slide 60: Illustration of ISP Parameter identification
Defined initial condition
simulation
Model parameter vector for
Pressure measurement
Parameter estimation by „Maximum-Likelihood“ principle
Predicted decrease of

## Slide 61: Implementation of the method in PARIO
Pressure measurement
The ISP method was made commercially available by a measuring instrument called PARIOTM, manufactured and traded by METER Group AG, Munich.