## Slide 1: Correction of siphoning error in processed historical rainfall intensity data, a case study of data measured by Hellmann-Fuess type rainfall recorder
Rácz Tibor MSc water eng.
EGU General Assembly 2022
Session CL5.1.5 – Downscaling: methods, applications and added value
Photo: Dr. Alfred Müller Catalogue

## Slide 2: Importance of the issue
Significant damage capacity of heavy rainfall caused floods.
Heaviest damage potential: short duration intensive rainfall.
Climate change issue: are the reference data accurate enough?

## Slide 3: Specific systematic error of siphoned rainfall recorders
Undermeasurement during siphonage
Reason: The emtying of the device takes time (10-30 s);
During the siphoning the rainfall recording is pending
But: The time of siphonage depends on the actual rainfall intensity
Correction: method of Luyckx and Berlamont, in the possession of registration banner
LUYCKX G., BERLAMONT J. (2002): Accuracy of siphoning rain gauges. In: Global Solutions for Urban Drainage. 2002. p. 1-12. https://doi.org/10.1061/40644(2002)251

## Slide 4: Correction of extracted rainfall intensity data
Extracted data: The highest sums (intensities) of a class of duration were collected during the processing of data
Correction: modified method of Luyckx and Berlamont
it Measured t minute intensity
icorr Adjusted t minute intensity
t Duration
ts,0 Siphonage time (no rain supply)
ho Initial water level
hs Measurable raindepth of the ribbon
Af Surface of collecting area
qs Drainage discharge of the device

## Slide 5
Correction of extracted data of Hellmann-Fuess recorder
Parameters of the device:
hs measurement capacity, width of the tape 10 mm
Af collection area of the device 20000 mm2
qs drainage discharge of the device 12500 mm3/s
ts,0 siphonage time (no rain supply) 18 s
Initial water level issue:
The initial water level in the measurement tank is a random variable
The waterlevel is somewhere between 0-10 mm, randomly.
ho can be estimated as expected value of a uniform distribution (E=hs/2=5 mm)

## Slide 6: Effect of correction on the rainfall’s PDF and IDF data
Adjusted data / Raw data, % |  |  |  |  |  | 
Average recurrency (years) | Exceedance Probability (%) | 5 min | 10 min | 20 min | 30 min | 60 min
1 | 99 | -2,3 | 3,7 | 0,0 | 1,9 | 0,0
2 | 50 | 5,4 | 3,0 | 2,2 | 1,3 | 1,1
10 | 10 | 9,0 | 5,1 | 3,6 | 2,7 | 1,7
50 | 2 | 11,1 | 7,4 | 4,9 | 3,9 | 2,4
100 | 1 | 11,9 | 8,6 | 5,4 | 4,3 | 2,7
Percentual differences between raw and adjusted IDF data
General Extreme Value PDF curves on raw and adjusted data
Highest corrections:
in the range of shorther duration than 30 min.
Significant corrections:
over 10 ys recurrency
Data: Hungarian Meteorological Service

## Slide 7: Summary and Conclusion
Correction of systematic errors is possible for the extracted data on a statistical basis.
The magnitude of the correction on short durations and rare intenisties is higher than 5%.
The effect of the correction is significant.
The improvement may affect the evaluation of climate change by refining the precipitation intensity reference data.
The repair will allow a reassessment of the suitability of existing drainage facilities.
Further information about the method:
https://www.researchgate.net/publication/354586738_On_the_correction_of_processed_historical_rainfall_data_of_siphoned_rainfall_recorders
https://www.researchgate.net/publication/351917435_Application_of_Correction_Procedures_for_Some_Systematic_Measurement_Errors_to_Rainfall_Intensity_Data_of_a_Rain_Gauge_in_Budapest

## Slide 8: Thank YOU for the kind attention!