## Slide 1
JAN-FRANTIŠEK KUBÁT · MICHAL VRÁNA · DAVID ZUMR · PETR KAVKA
A LASER DIFFRACTOMETRY TECHNIQUE FOR DETERMINING THE SOIL WATER STABLE AGGREGATES INDEX

## Slide 2
WATER STABLE AGGREGATES INDEX
WSAi
WHY
Gives us a quick indication of the soil condition
AIM
Develop a technique that is comparable to the commonly used wet sieving technique
Fast, reliable and simple to perform
METHODS
Six sampling sites
Eijkelkamp wet sieving apparatus – wet-sieving (WS technique)
Mastersizer 3000 – sonication and laser diffractometry (SLD technique)

## Slide 3
INTRODUCTION OF SLD TECHNIQUE
SLD - SONICATION AND LASER DIFFRACTOMETRY
PRINCIPLE
Comparison of two particle size distributions
Kemper and Rosenau
Kubát and Vrána (simplified eq. – full eq. here)

## Slide 4
RESULTS
SETTINGS
WORKFLOW
PRINCIPLE
tap to navigate
A Laser Diffractometry Technique for Determining the Soil Water Stable Aggregates Index
PRINCIPLE
JAN-FRANTIŠEK KUBÁT · MICHAL VRÁNA · DAVID ZUMR · PETR KAVKA
NAVIGATION MAP

## Slide 5
WORKFLOW
PRINCIPLE
SETTINGS
RESULTS
NAVIGATION
PRINCIPLE
SRsol,n is the volume residue in the n-th size class in the dispersion solution (%)
SRw,n is the residue in the n-th size class in water (%)
WSAiSLD,soil is the water stable aggregates index of the soil (%)
n is the number of size classes (80 in this case)
Comparison of particle size distributions
A Laser Diffractometry Technique for Determining the Soil Water Stable Aggregates Index

## Slide 6
WORKFLOW
PRINCIPLE
SETTINGS
RESULTS
NAVIGATION
A Laser Diffractometry Technique for Determining the Soil Water Stable Aggregates Index
WORKFLOW

## Slide 7
WORKFLOW
PRINCIPLE
SETTINGS
RESULTS
NAVIGATION
Parameter |  | Distilled water | Dispersion solution
Soil properties | Refractive index | 1.457 | 
 | Absorption index | 0.01 | 
 | Particle density (g·cm-3) | 2.64 | 
 | Sample amount (g) | 0.4 | 
Properties of the solution | Refractive index | 1.33 | 
Hydro LV properties | Sonicating level (%) | 0 | 100
 | Sonicating time (s) | 0 | 70
 | Stirrer/pump speed (rpm) | 2500 | 
Measurement Properties | Red-laser duration (s) | 10 | 
 | Blue-laser duration (s) |  | 
 | Repetitions | 5 | 
 | Weighted residua (%) | < 1 | 
Total operating time (s) |  | 100 | 170
Total measurement time (s) |  | 100 | 
Mastersizer 3000
Instrument settings during the measurement of the particle size distribution
A Laser Diffractometry Technique for Determining the Soil Water Stable Aggregates Index
SETTINGS

## Slide 8
WORKFLOW
PRINCIPLE
SETTINGS
RESULTS
NAVIGATION
RMSE – 6.7 %
ANOVA p-value – 0.11
A Laser Diffractometry Technique for Determining the Soil Water Stable Aggregates Index
RESULTS