## Slide 1: Development of a Dual‑Chamber Atmospheric Simulation System for Bioaerosol Research: Size‑Dependent Analysis and Surface Interaction Studies
Fumito Maruyama and So Fujiyoshi
Microbial Genomics and Ecology,
Hiroshima University,
The IDEC Institute, Japan
EGU25-1931 (AS4.17):
PICO spot 3 at 10:45-10:47, Vienna, May 2nd, 2025

## Slide 2
Overview: Development of a Dual-Chamber Atmospheric Simulation System for Bioaerosol Research
Custom dual bioaerosol chamber (BSL2, UV, Temp control)
Outdoor comparison of two real-time bioaerosol counters
Indoor comparison of two real-time bioaerosol counters
Case Study: Bioaerosol Measurement at a University Hospital
* Dual-chamber system enabled stable, reproducible bioaerosol measurements.
* Field tests showed particle counts agreement; Kanomax detected higher biological variability.
* Hospital study linked larger bioaerosol peaks to human activity; smaller particles dominated by background aerosols.
Airborne microorganisms (bioaerosols) play essential roles in global biogeochemical cycles and ecosystem dynamics
Integrated lab–field approach offers an unprecedented platform for bioaerosol research, significantly advancing understanding of microbial aerosol dynamics and their environmental and health impacts.

## Slide 3
Viewing 11:07-12:30 at PICO3.1
fumito@hiroshima-u.ac.jp
EGU General Assembly 2025, Vienna, Austria
Google:
MGE, HIROSHIMA
Notes: Regarding the EGU General Assembly 2025, we hereby inform you that your following abstract has been re-scheduled based on the decision of the respective convener: EGU25-1931 &nbsp;Development of a Dual-Chamber Atmospheric Simulation System for Bioaerosol Research: Size-Dependent Analysis and Surface Interaction Studies by Fumito Maruyama and So Fujiyoshi Your contribution is now scheduled as PICO presentation in session AS4.17 - Characterization, processes, and surface interactions of the atmospheric microbiome. The oral presentation of your PICO, 2-minute summary, takes place at PICO spot 3 on Friday, 2 May 2025, 10:45. The PICO viewing time will be Friday, 2 May 2025, 11:07-12:30 at PICO screen PICO3.1. For details of your presentation as well as of the corresponding session programme please see: https://jpn01.safelinks.protection.outlook.com/?url=https%3A%2F%2Fmeetingorganizer.copernicus.org%2FEGU25%2Fsession%2F53552&data=05%7C02%7Cfumito%40hiroshima-u.ac.jp%7C5ae0b8be09744f31d2ec08dd817f7763%7Cc40454ddb2634926868d8e12640d3750%7C1%7C0%7C638809106672334840%7CUnknown%7CTWFpbGZsb3d8eyJFbXB0eU1hcGkiOnRydWUsIlYiOiIwLjAuMDAwMCIsIlAiOiJXaW4zMiIsIkFOIjoiTWFpbCIsIldUIjoyfQ%3D%3D%7C0%7C%7C%7C&sdata=Ba%2BVgoS2vRPDINeCqb0KvSSfnNnhrr1eWWCyxJKkoI4%3D&reserved=0. We thank you very much for your cooperation and we are looking forward to seeing you at the EGU General Assembly 2025. In case any questions arise, please do not hesitate to contact us. Kind regards, The Copernicus Meetings team egu25@copernicus.org
https://meetingorganizer.copernicus.org/EGU25/EGU25-1931.html

## Slide 4: Overview

## Slide 5
Development of a Dual-Chamber Atmospheric Simulation System for Bioaerosol Research
Airborne microorganisms, or bioaerosols, play essential roles in global biogeochemical cycles and ecosystem dynamics. To study their physicochemical properties and survival mechanisms, a novel dual-chamber atmospheric simulation system has been developed. This innovative platform allows detailed investigation under controlled conditions, combining advanced technology and environmental precision to unlock new insights into bioaerosol behaviour.
This research offers crucial advancements for environmental microbiology and atmospheric sciences, supporting comprehensive exploration of microbial transport and surface interactions with unprecedented accuracy.
FM
Overview
NOT real!

## Slide 6
System Architecture and Clean Room Installation
Dual Stainless Steel Chambers
Designed with two interconnected stainless steel chambers to ensure precise environmental control. Chambers minimize microbial adhesion and prevent electrostatic losses, enhancing experiment reliability.
BSL-2 Class 100 Clean Room Installation
Installed in a biocontainment Level 2 clean room certified to Class 100 standards, guaranteeing a highly sterile and safe environment essential for accurate bioaerosol experiments.
Overview

## Slide 7
Environmental Controls and Interconnected Chamber Design
UV Irradiation and Temperature Control
Each chamber includes dedicated UV irradiation and temperature control systems, allowing simulation of environmental stressors such as radiation and thermal changes.
Interconnected Chambers with Dampers
Dampers permit independent environmental regulation in each chamber, enabling bioaerosol generation in one while exposure experiments occur in the other.
Surface Treatment Optimization
Advanced surface treatments were applied to reduce microbial adhesion, ensuring accurate aerosol behaviour and minimal experimental artefacts.
NOT real!
Overview

## Slide 8
Simultaneous Physical and Biological Monitoring
Real-Time Particle Counting
Continuous monitoring of aerosol particle size and concentration enables observation of dynamic changes during experiments.
SEM-EDS Morphology and Composition
Scanning electron microscopy with energy-dispersive X-ray spectroscopy reveals detailed particle morphology and elemental composition.
Biological Activity Assessment
Concurrent evaluation of microbial viability supports understanding of survival strategies under various environmental stresses.
Overview
NOT real!

## Slide 9
Size-Dependent Effects and Surface Interaction Studies
Size Segregation Capability
The system segregates particles by size, allowing targeted analysis of size-dependent disinfectant efficacy and environmental stress effects.
Surface Interaction Optimisation
Surface properties of the chambers are fine-tuned to study bioaerosol adhesion and retention, critical for understanding microbial transport mechanisms.
Overview

## Slide 10
Validation and Experimental Applications
The system’s effectiveness was validated through rigorous experiments with various environmental microorganisms. Results demonstrate stable conditions and precise control enabling reproducible measurement of bioaerosol physical characteristics and biological activity.
Such validation confirms this platform's utility for a broad range of bioaerosol research, including pathogen survival, environmental stress response, and aerosol transport studies.
Overview

## Slide 11
Current Research Focus: Airborne Microorganism Transport
Bioaerosol Transport Processes
Investigating how microorganisms travel through the atmosphere under varying environmental conditions.
Atmospheric Component Interactions
Studying interactions between bioaerosols and atmospheric particles such as dust, pollutants, and water vapour.
Environmental Impact Assessment
Determining implications for global biogeochemical cycling and ecosystem dynamics influenced by airborne microbes.
Overview

## Slide 12
Significance and Future Directions
This dual-chamber system provides an unprecedented platform for comprehensive bioaerosol research, integrating physical characterisation with biological analysis. It paves the way for enhanced understanding of microbial ecology in the atmosphere and informs strategies for infection control and environmental management.
Future developments may include automation, integration with molecular techniques, and expanded environmental simulation capabilities to further advance bioaerosol science.
Overview

## Slide 13: Bioaerosol chamber

## Slide 14
FM
Chamber
Three rooms
BSL2, C100

## Slide 15
23.2L x 2 rooms
Temperature
Pressure
UV light
Toward precise evaluation of dynamics of aerosol
Photo: Taken by presenter
https://doi.org/10.1016/j.jphotobiol.2020.111899.
Chamber
Bioaerosol Research Setup and Testing
Notes: 23.2L x 2

## Slide 16
FM
Chamber
Chamber characteristics

## Slide 17
Syringe Pump Configuration
Set Syringe Diameter
Input syringe diameter and wait 2 seconds to confirm.
Select Flow Rate
Choose mL/min rate and enter desired value.
Setup Procedure
Attach tubing, draw liquid, purge air, and connect to system.
Maintain Pump Pressure
Keep pressure at 0.1 MPa during operation for reliability.
Chamber

## Slide 18
Substances Added to Chamber
MQ Water
Pure H2O used as liquid base
70% Ethanol
Organic solvent: C2H6O
Phosphate Buffered Saline
PBS solution for stability
Chamber

## Slide 19
Nanoparticle Size Distribution of PBS Solution
Optimal PBS particle size measured at 181.24 nm with clear generation and decay phases.
Glycerol solution plus SLAGS aerosolizer produces 55.49 nm nanoparticles (not shown in graphs).
Chamber
PROBLEM happened!!
Unknown PM appeared

## Slide 20
Chamber
Particle Analysis in Aerosol Chamber
SEM-EDX used for elemental analysis
4 locations examined inside chamber
Purpose
Identification of particle composition

## Slide 21
Source Side Tube Mouth Analysis
Particles uniform morphology; stainless steel alloy composition suspected
Chamber
Solenoid valve in Aerosol Chamber caused PM!
HEPA capsule was implemented

## Slide 22
Bubble Coverage Control in Slag Plate
Bubble Coverage Adjustment
Adjust liquid amount to cover slag plate partially or fully with foam.
Chamber

## Slide 23
Aerosolization of Injection Water at 8.5 mL/min
Generation side
Key Observations
Particle transfer confirmed upon damper opening.
Syringe pump adjustment needed for stable operation.
Single operator struggles with detailed operation logging.
Cleaning increased particle count on generation side only.
Systematic operation control is essential.
Chamber
Recovery side

## Slide 24: Outdoor comparison

## Slide 25
Outdoor comparison of two real-time bioaerosol counters
Analysis of total and biological particle counts using KANOMAX and BioTrak instruments in 2024.
Outdoor

## Slide 26
Initial 10-Minute Continuous Measurement
Total particle trends
Consistent fluctuations between devices
Particle count differences
KANOMAX detects fewer particles
Biological particle timing
Minimum occurs at different times
Outdoor

## Slide 27
2024.10.25
KANOMAX (2024/10/2-21)
2.83 L/min
BioTrak (2024/10/2-21)
28.3 L/min
φ2_3 µm
φ1_3 µm
Bio
Similar dynamics
Y軸の値に関してはKANOMAX＞BioTrak
Total consistent
Latest 10-Minute Measurement Overview
Particle count agreement
Total particles mostly match between devices
Higher counts in KANOMAX
KANOMAX readings slightly exceed BioTrak
Similar biological particle trends
Variations align closely
Outdoor

## Slide 28
R²: 0.628
R²: 0.557
R²: 0.54
R²: 0.743
Correlation Coefficients Overview
R² range: 0.54 to 0.74
Indicating moderate to strong linear relationships
Highest correlation at 3-5 µm
R² = 0.743, strongest explanatory power
All size ranges show positive correlation
r values above 0.7 indicate robustness
Outdoor
Notes: サンプラー入り口あっためる。
コリオリス->BioSpot->FlowCam

## Slide 29
Particle Size Range Correlations
Size Range (µm)
Correlation (r)
R²
0.5-1
0.793
0.628
1-2
0.747
0.557
2-3
0.735
0.540
3-5
0.862
0.743
Stronger correlation with increasing particle size
Outdoor

## Slide 30
Time-Variation Particle Count Comparison
Total Particles (φ1_3 µm)
Nearly identical variation trends between instruments
Biological Particles (φ2_3 µm)
KANOMAX consistently higher, notably 10/14–16
Outdoor

## Slide 31
Correlation Between Instruments
Total Particles
Correlation coefficient R=0.945, highly significant
Biological Particles
Moderate correlation R=0.659, p-values near zero
Outdoor

## Slide 32
statistic | BioTrak calactV1 | KANOMAX Bio2_3_perL
①mean | 2.954432 | 14.2853619
②median | 1.413428 | 8.975265
③sd | 3.280313 | 15.4469754
④q1 | 0 | 5.8303887
⑤q3 | 4.240283 | 15.3003534
⑥min | 0 | 0.5300353
⑦max | 25.441696 | 133.9222615
⑧n | 2803 | 2020
⑨Wilcoxon p-value（中央値比較） | 0.000000 | NA
⑩KS-test p-value （分布の形状比較） | 0.000000 | NA
Biological Particle Distribution Comparison
KANOMAX higher mean
About 4.8 times BioTrak average
Greater variability
KANOMAX shows larger data spread
Significant distribution difference
Statistical tests show p<0.001
Outdoor

## Slide 33
Summary and Implications
Measurement Agreement
KANOMAX and BioTrak generally align in total counts
Biological Particle Variance
KANOMAX reports higher bio-particles, wider range
Correlation Strength
Strong positive correlations across size ranges
Future Work
Refine calibration for biological particle accuracy
NOT real!
Outdoor

## Slide 34: Indoor comparison

## Slide 35: Comparison of two real-time bioaerosol counters
Dec 13th 00:25 (Thu) ~ Dec 16th 11:05 (Mon)
1 min measure and 9 min interval
28.3 L/min
Fumito MARYAMA and So FUJIYOSHI
Dec29th, 2024
Indoor
Notes: Indoor validation: Comparison of two real-time bioaerosol counters

## Slide 36
Simultaneous monitoring with prototype device from Kanomax
Biotrak is much more noisy
Kanomax measures more large particles. (3-5, 5<)
Indoor

## Slide 37
Screen
Window
Window
AC
Veranda
Window
Window
Window
Burner
Fan
Corridor
Corridor
Restroom
Basin
Laundry
Bathroom
Desk
Sofa
Sofa
Table
Cooker
Fridge
Airdog
Roomba
Micro
Dish washer
Sink
Table
Food dryer
Kanomax
BioTrak
Stairs to
Entrance
Stairs to
loft
Monitoring site
Set at a volunteer’s living room of an apartment in Higashi-Hiroshima, Hiroshima, Japan
Indoor

## Slide 38
2024/12/16 11:05
2024/12/13 19:05
102
101
100
103
102
101
100
103
104
Sat
Sun
Mon
Total PM
Viable PM
Total 0.5, 0.7, 1.0 similar, 3.0 unique, and 5.0, 10.0 similar dynamics
Viable 0.5, 0.7 none, 1.0 unique, 3.0, 5.0, 10.0 similar dynamics
105
BioTrak
Indoor

## Slide 39
2024/12/16 11:05
2024/12/13 19:05
0
250
0
80,000
Sat
Sun
Mon
Total PM
Viable PM
Total 0.5 unique, 0.7 and 1.0 similar dynamics; PM2.5 and PM10 are not useful for the biological air quality
Total and Viable are not correlated, meaning Total is not Viable alternative
BioTrak
Indoor

## Slide 40
2024/12/16 11:05
2024/12/13 19:05
Sat
Sun
Mon
Total
Viable
Total 5.0 appears more related to Viable 1.0
Viable 3.0 are more related to human activities
Log Total
Log Viable
BioTrak
Indoor

## Slide 41
2024/12/17 7:31
2024/12/13 20:31
102
101
100
103
Sat
Sun
Mon
Total PM
Viable PM
102
101
100
103
104
105
Tue
Kanomax
Total 2.0 is similar to 0.5 and 1.0 and same T and V size sometime similar dynamics
Viable 0.5,1.0,2.0 similar, 3.0 and 5.0 similar, dynamic change in abundance
104
Indoor

## Slide 42
Biotrak T
Kanomax T
BioTrak and Kanomax smilar and Total 2.0 is between 1.0 and 3.0.
BioTrak 5.0 and 10.0 is clear but Kanomax 5.0> is intermittent
Indoor

## Slide 43
Biotrak V
Kanomax V
102
101
100
103
104
102
101
100
103
Large V sometimes higher than smaller in Kanomax
Human activities correlated large V
Small V dynamics are similar
Kanomax provide 0.5 and 10 dynamics
Indoor

## Slide 44: Next Steps
More finer scale fractionation (particle size distribution)
More comparison with other devices
Detailed correlation of human activities or behavior with total and viable aerosol (bath, shower, fermented food, etc)
Difference between Kanomax and BioTrak
Indoor

## Slide 45: Case study

## Slide 46
Bioaerosol Measurement
at a University Hospital
This presentation shares results from bioaerosol measurements conducted at a University Hospital in April 2025, analyzing particle distribution and human traffic correlation.
NOT real!
Case Study
Notes: Case Study: Bioaerosol Measurement at a University Hospital

## Slide 47
Sensing Principle
Light Scattering
All particles produce scattered light when illuminated.
Fluorescence Detection
Bio-particles emit fluorescence when illuminated.
Particle Classification
Simultaneous detection of scattered light and fluorescence identifies bio-particles.
Case Study

## Slide 48
Measurement Setup
1
Date and Time
April 8-10, 2025. Two sessions: 17:51-10:30 and 17:53-3:32.
2
Location
Corridor in front of negative pressure room, 8th floor, West Wing.
3
Equipment
Prototype bioaerosol sensor calibrated with standard particles.
4
Protocol
999 measurements per minute for 16 hours 39 minutes.
NOT real!
Real
Case Study

## Slide 49
Traffic Patterns (number of person/hour)
Traffic peaked during morning, noon, and evening hours. Most entries were recorded as "passing through." Nighttime showed minimal traffic.
Case Study

## Slide 50
Particle Size Distribution (Total, Bio, Ratio)
0.5-1.0
1.0-2.0
2.0-3.0
3.0-5.0
> 5.0
Case Study

## Slide 51
Case Study

## Slide 52
Particle Size Distribution
1
0.5-1μm
Highest concentration, likely dust particles
2
1-2μm
High concentration throughout measurement period
3
2-3μm
Moderate correlation with human traffic
4
3-5μm
Strong correlation with human presence
5
>5μm
Highest human-derived correlation
Case Study

## Slide 53
Measurement Results
Small Particles (0.5-2μm)
High values throughout all time periods
Peak concentrations around midnight
No correlation with human traffic
Likely dust or environmental particles
Large Particles (>2μm)
Lower values during nighttime
Strong correlation with human traffic
Likely human-derived bioparticles
Less affected by air currents
Case Study

## Slide 54
Correlation Analysis
Small Particles
Negative correlation with traffic. Possibly due to increased airflow to negative pressure room.
1
Large Particles
Positive correlation with traffic. Likely human-derived bioaerosols.
2
Door Opening
May reduce small particle concentration through airflow to negative pressure room.
Particle Resistance
Larger particles resist airflow effects due to greater mass.
4
Case Study

## Slide 55
Key Findings
Time Patterns
Particle counts increased during morning, noon, and evening. Coincided with higher traffic periods.
Size Differentiation
Small particles (0.5-2μm): Environmental origin. Large particles (>2μm): Human-derived.
Traffic Correlation
Negative correlation for small particles. Positive correlation for large particles.
Airflow Hypothesis
Door opening frequency may affect small particle concentration through negative pressure effects.
NOT real!
Case Study

## Slide 56
Resistome in Bioaerosol
Pierre AMATO
Xavier RODO
Fumito MARUYAMA
Acknowledgement

## Slide 57
Research Grant and Collaboration Support
Japan Prize Heisei Memorial Research Grant
Kurita Water and Environment Foundation
Kyoto University & DAIKIN Industries joint grant
Collaboration with CHOBE, Hiroshima University
Japan Society for the Promotion of Science (KAKENHI)
Japan Agency for Medical Research & Development (AMED)
Japan-Slovenia Research cooperative programs
Asahi glass foundation
Institute of Nature and Environmental Technology, Kanazawa University
Acknowledgement