## Slide 1
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface
Nicolò Tonolo1, Sergio Teggi2, Simona Berardi3, Maria Paola Bogliolo3, Iason Verginelli1
1 Laboratory of Environmental Engineering, Department of Civil Engineering and Computer Science Engineering, University of Rome Tor Vergata
2 “Enzo Ferrari” Department of Engineering, University of Modena and Reggio Emilia
3 Department of Technological Innovations and Safety of Plants, Products and Anthropic Settlements- National Institute for Insurance against Accidents at Work (INAIL)
Session BG2.10, EGU 2026

## Slide 2
Flux Chamber Description
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
1
Stainless Steel Cylindrical Flux Chamber (V ~ 6 L)
Photoionization Detector (PID) sensor for VOCs measurements
Non-Dispersive Infrared (NDIR) sensor for CO2 measurements
LoRaWAN transmission system for remote connectivity to IoT clouds
Photovoltaic panel for energy self-sufficiency
PID sensor
NDIR sensor
Vent
External Hardware connector
Stainless Steel Chamber

## Slide 3
Flux Chamber Description
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
1
LoRaWAN transmission system
Photovoltaic Panel
Stainless Steel Cylindrical Flux Chamber (V ~ 6 L)
Photoionization Detector (PID) sensor for VOCs measurements
Non-Dispersive Infrared (NDIR) sensor for CO2 measurements
LoRaWAN transmission system for remote connectivity to IoT clouds
Photovoltaic panel for energy self-sufficiency

## Slide 4
Principle of Operation
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
2
Measured flux transitted to the IoT cloud
Autonomous measurement every 20 min
Concentration increase measured by the internal sensors
Measurement Phase (10 min)
Activation of two micro-pumps for internal air exchange. Generation of optimal conditions for a new measurement phase
Air Exchange Phase (10 min)

## Slide 5
Laboratory Emission Test
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
3
High stability in consecutive measurements
Elevated linearity of regressions for flux calculation

## Slide 6
Monitoring Campaign at a Real Site
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
4
Trend consistent with day/night temperature variations
Emission reduction in rainy days
12 days of continuous measurements
Collected ~ 2000 VOCs and CO2 flux measurements

## Slide 7
The monitoring of Volatile Organic Compound (VOC) emissions and their natural attenuation from the subsurface is typically conducted by specialized operators through the collection, transport, and laboratory analysis of samples
Monitoring carried out on a seasonal basis, without temporal continuity
High operational costs
Limited data availability (~ 4 measurements/year)
VOC and CO2 Monitoring at Contaminated Sites
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
5
Groundwater
Unsaturated Soil
Fuel Spill
VOCs Diffusion and Attenuation
CO2 production

## Slide 8
Temperature
Humidity
Precipitation
Variation in contamination
Subsurface Heterogeneity
Need for continuous, low-cost monitoring of subsurface emissions
VOC and CO2 Monitoring at Contaminated Sites
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
5
Influencing Parameters
Groundwater
Unsaturated Soil
Fuel Spill
VOCs Diffusion and Attenuation
CO2 production

## Slide 9
Component | Quantity | Model | Unitary Price
Microcontroller | 1 | Arduino Uno R4 wifi | 30 €
Analogue Digital Converter (ADC) | 1 | ADS1115 | 20 €
Temperature, Humidity and Pressure Sensor | 1 | BME280 | 20 €
VOC Sensor | 1 | PID-AH5 | 500 €
VOC Sensor | 1 | PID-A15 | 450 €
CO2 Sensor | 1 | SCD41 | 50 €
LoRaWan Transmission Module | 1 | Nucleo WL55JC1 | 53 €
Lead-Acid Battery | 1 | Pro-elec 12V, 16.5Ah | 35 €
H Bridge | 1 | DRI0042 | 20 €
Micro Air Pump | 2 | BTX-MVP-B1M | 300 €
SD Module | 1 | DFR0229 | 6 €
SD Card | 1 | Micro SD kingstone, 32GB | 8 €
Electronic Hardware Housing | 1 | ABB (310x240x110 mm) | 35€
Sensor Housing | 1 | 3D Printed | 3 €
Stainless Steel Chamber | 1 | Self Designed | 450 €
MPPT Regulator | 1 | MPPT | 20 €
Photovoltaic Panel | 1 | Enjoy Solar 100W | 100 €
 |  | Total | 2400 €
Flux Chamber Description
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
6

## Slide 10
Calibration Test
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
7
Test Setup
Calibration conducted comparing sensor signal with concentration measurements of a commercial VOCs analyzer
1,2,8 L/min Nitrogen Fluxes for generating concentration fluctuations with different gradients during the test.
Calibration tests performed with 6 different organic compounds

## Slide 11
Calibration Test
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
7
CV < 20% of calibration parameters
Results independent from the calibration compound

## Slide 12
25 cm
Artificially contaminated soil
Emission test from an artificially contaminated soil
Evaluating VOCs natural attenuation processes in the subsurface
Test replicated with two different fuel mixtures (i.e. gasoline and diesel)
Continuous monitoring over a two month period
Laboratory Emission Test
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
8
VOCs Diffusion and Biodegradation

## Slide 13
Laboratory Emission Test
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
8
Emission test (1–2 months) to evaluate the device’s performance in continuously monitoring the natural attenuation of two fuel mixtures

## Slide 14
Monitoring Campaign at a Real Site
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
9
Measurements representative of real site-specific atmospheric conditions
No significant overheating within the chamber volume

## Slide 15
An Innovative IoT-Based Automated Flux Chamber for Continuous Monitoring of CO2 and VOC Emissions from the Subsurface (Session BG2.10)
EGU26 Abstract
Reference
10