## Slide 1: Quantifying Biases in Open-path Eddy Covariance CO2 Flux Measurements Caused by Spectroscopic Effects in Broadband Non-dispersive Infrared Gas Analyzers
Ivan Bogoev
Campbell Scientific, UT, USA
EMS 2022, Bonn
Notes: I would like to thank the organizers for the opportunity to present.

## Slide 2: Motivation: controversial findings
Ma, J. et al. (2014). A downward CO2 flux seems to have nowhere to go. Biogeosciences, 11: 6251-6262.
Schlesinger, W. (2016). An evaluation of abiotic carbon sinks in deserts. Global Change Biol., doi: 10.1111/gcb.13336
Is the desert a carbon sink? Where does the carbon go?

## Slide 3: Motivation: unreasonable CO2 uptake

## Slide 4: Motivation: unreasonable CO2 uptake

## Slide 5: Motivation: unreasonable CO2 uptake
64 total sites
37 sites
show bias

## Slide 6: Motivation: poor comparisons between open- and closed-path analyzers

## Slide 7: Biases in open-path sensor is driven by sensible heat flux

## Slide 8: Background: broadband non-dispersive infrared spectroscopy
A = Absorbed light energy in the spectral interval
N = Number of absorption lines in the spectral interval
ν = Wave number of the individual spectral line
c = Density of the absorbing gas
L = Path length
Si = Strength of the individual line
αi = Half-width of the individual line
Si = f(T) αi = f(T,P,)
T = Air temperature, P = barometric pressure
A = f(c,T)

## Slide 9: Non-dispersive infrared spectroscopy
Non-dispersive: interference filters
Dispersive
After Infratek Gmbh

## Slide 10: Methods: HITRAN simulations
600 ppm 244 K, 29.95 mmol m-3
700 ppm 285 K, 29.95 mmol m-3
700 ppm 326 K, 29.95 mmol m-3
710 ppm 285 K, 30.35 mmol m-3
Same CO2 density at different temperatures

## Slide 11: Collision broadening: weak lines
326 K
285 K
244 K
Doppler broadening of weak lines where line strength and line width increase with increase in temperature

## Slide 12: Collision broadening: strong lines
Doppler broadening of strong lines where line strength and line width decrease with increase in temperature
244 K
285 K
326 K

## Slide 13: Methods: HITRAN integrated spectra
A
B
2392 cm-1
2309 cm-1
2316 cm-1
Case A
Case B
Optimal
Accounting for filter tolerances
Δ
2392 cm-1

## Slide 14: Results: HITRAN integrated spectra
Case A
Case B
No temperature dependance
Δ10 ppm CO2
Δ10 ppm CO2
Δ82 K
Temperature sensitivity approx. 0.25 ppm K-1
or 0.0098 mmol m-3 K-1
Warmer air will cause underestimation of CO2

## Slide 15: Methods: synthetic calibration HITRAN
CO2 concentrations: 0 to 1000 ppm
100 ppm
200 ppm
300 ppm
1000 ppm
400 ppm
T [K]
316 K
295 K
276 K
Temperatures: 276, 295 and 316 K
Pressure: 101.4 kPa
Generated 33 CO2 absorption spectra for spectral range case B

## Slide 16: Results: synthetic calibration residuals
]
Standard - measured
200 ppm
300 ppm
1000 ppm
400 ppm
316 K
295 K
276 K
underestimation
100 ppm
overestimation

## Slide 17: Discussion
μ

## Slide 18: Conclusions
There is an optimal spectral band in the CO2 absorption spectrum where line-broadening temperature effects are balanced
When the optimal CO2 absorption band is properly sampled by the interference filter the temperature effects are negligible. Case A
A manufacturing tolerance of the filter pass-band can increase the sensitivity of the analyzer to air temperature changes and if not corrected can cause a systematic error in measured CO2 concentrations. Case B
2309 cm-1
2316 cm-1
Case A
Case B
2392 cm-1

## Slide 19: Conclusions
Open-path broad-band non-dispersive infrared CO2 gas analyzers need a correction for temperature related spectroscopic effects
Accurate and fast-response air temperature measurements in the optical path of the analyzer are needed to properly apply the correction
A = f(c,T)
T air

## Slide 20: References
Wang et al. (2017) A meta-analysis of open-path eddy covariance observations of apparent CO2 flux in cold conditions in FLUXNET
J. Atmos. Ocean. Technol., 34
Wang et al. (2017) Performance evaluation of an integrated open-path eddy covariance system in a cold desert environment
J. Atmos. Ocean. Technol., 33
Ono et al. (2008) Apparent downward CO2 flux observed with open-path eddy covariance over a non-vegetated surface
Theor. Appl. Climatol., 92, pp. 195-208
Russell et al. (2019) Adjustment of CO2 flux measurements due to the bias in the EC150 infrared gas analyzer
Agric. For. Meteorol., 276 Article 107593