## Slide 1: A Doppler Lidar Technique for Monitoring the „Whole“ Boundary Layer UP 1.5 Hans-Jürgen Kirtzel, Dr. Gerhard Peters Metek GmbH

## Slide 2: A Doppler Lidar Technique for Monitoring the „Whole“ Boundary Layer Hans-Jürgen Kirtzel, Dr. Gerhard Peters Metek GmbH
Monitoring: Routine unattended observation with automatic online QA / reliability proof of data
Boundary Layer: No unique definition, can be identified by profiles of wind, temperature, turbulence, aerosol, or by range availability of remote profilers
Whole: From near ground heights (≈10 m, height of small masts) up to geostrophic wind field
Our Target: Wind and turbulence, Lidar

## Slide 3: Motivation
How much information yields a wind profile?
The strongest gradients are typically observed near the surface.

## Slide 4: Motivation
How much information yields a wind profile?
The strongest gradients are typically observed near the surface.
70 m -
500 m -
1000 m -
1500 m -

## Slide 5: Motivation
Aviation: Flight safety
Airport management
Reduction of environmental impact
Environmental applications: Air quality
Dispersion of released materials
Cold outflows
Heated metropols
Engineering tasks: Wind and ghust loads
Wind energy (ultra high wind turbines, > 250 m, sail systems)
Weather forecast: Modelling
tbc
70 m -
500 m -
1000 m -
1500 m -

## Slide 6: Doppler Lidar System s (Based on aerosol scattering)
Pulsed Lidars: Ranging by simple time counting High peak pulse power
Unambiguos range determination Limited range resolution
High SNR for a given mean power Limited minimum measuring height
CW Lidars: Ranging by optical focal setting Data affected by low hanging clouds
Relatively small emission peak power Unknown sign of radial wind
Fine range resolution in low ranges Blind zone around zero
Low minimum measuring height Limited to max. ranges of 200/300 m
Consecutive measurements in different heights
FM-CW Lidars: Ranging by optical focal setting Data affected by low hanging clouds
Relatively small emission peak power Unknown sign of radial wind
Fine range resolution in low ranges Blind zone around zero
Very low minimum measuring height Limited to approx. 200/300 m ranges
Determination of effective Consecutive measurements in different heights
measuring height

## Slide 7: FM CW Doppler Lidar Wind Ranger 100/200
Wind Ranger
100/200
VAD scans at 80° elevation at a rate of 0.5 … 2 rps
Sample rate 100 Hz
No 180° ambiguity, no blind zero zone
Up to 20 measuring heights available (typ. 6 … 10)
Determination of the effective measuring range
Measuring heights 7 … 200 m, range resolution 0.05 … 32 m
Narrow scan angle 10 °, in favor of vertical wind measurements
Compact and light weight, 30 cm x 40 cm x 50 cm w/o feet, 45 kg

## Slide 8: CW Doppler Lidar Systems Ranging and Range Resolution
3 “ lens diameter
Height / Lens diameter | 2“ lens | 3“ lens
10 m | 0,16 m | 0,08 m
20 m | 0,64 m | 0,32
100 m | 16 m | 8
200 m | 64 m /NA | 32 m

## Slide 9: FM CW Doppler Lidar Wind Ranger 100/200 Principle of triangular frequency modulation
Time
Optical Frequency
Time
Optical Frequency
Time
Optical Frequency
Case 1 velocity = 0:
Beat frequency at receiver depends only on distance to target
|
Case 2 velocity > 0:
Beat frequency at receiver depends on
distance to target and Doppler shift
|
Case 3 velocity < 0:
Beat frequency at receiver depends on
distance to target and Doppler shift
|
|
, i.e. height
, i.e. wind component
, i.e. height
fu
fd
fu
fu
fd
fd
fu and fd are the frequency shift between transmitted and received signal

## Slide 10: FM CW Doppler Lidar Wind Ranger 100/200 11.08.2022 00:00 – 23:50

## Slide 11: FM CW Doppler Lidar Wind Ranger 100/200 Discrimination between hard target and atmospheric signal
Internal reflection, ≈ 2 m distance
Wind signal, 10 m distance
Reflection from hard target, 20 m distance

## Slide 12: Pulsed Doppler Lidar Systems Variants
Available for different applications and corresponding ranges and heights
Short range (N beams or 2D scanner) | Range 40/60 … 200/350 m Compact (60 x 60 x 60 cm) Light weight (50 … 80 kg) Low power (< 170 W + 170 W) | Wind siting Urban areas Dispersion Wind systems in low heights
Long range (2D Scanner) | Range 70/100 … 2500/10000 m Height 70/100 … 2500 (PBL) Partly compact (60 x 60 x 60 cm / 90 x 90 x 90 cm) Partly light weight (80 kg/ >200 kg) Partly low power (< 170 W + 170 W / 1600 W) | Wind field investigation under low elevation Airports (glide pathwind measurements) Investigation of wind systems

## Slide 13: Pulsed Doppler Lidar Systems Range Resolution vs Velocity Resolution
(≈ 115 m2/s)
The relation between range resolution and velocity resolution is given by

## Slide 14: Pulsed Doppler Lidar Wind Ranger BL
Wind Ranger
BL
VAD scans at 80° elevation at a rate of 0.1 rps
Pulse rate 10 kHz, pulse length 320 ns, i.e. 100 m
Up to 100 measuring heights available
Measuring heights 80 … 2500 m
Narrow scan angle 10 °, in favor of vertical wind measurements
Compact and light weight, 30 cm x 40 cm x 50 cm w/o feet, 45 kg

## Slide 15: Pulsed Doppler Lidar Wind Ranger BL, 24.03.2022
Wind Ranger BL Messungen, Wind vectors, D-A-CH 2022 Leipzig,
Ideal conditions (red lines indicate 500/1000/1500 m)

## Slide 16: Pulsed Doppler Lidar Wind Ranger BL, 24.03.2022
Wind Ranger BL Messungen, Wind vectors, D-A-CH 2022 Leipzig,
Ideal conditions (red lines indicate 500/1000/1500 m)
Red area indicate blind zone up to 85 m due to pulse length of Lidar system (320 ns)

## Slide 17: Pulsed Doppler Lidar Wind Ranger BL, 24.03.2022
Wind Ranger BL Messungen, std.dev. of vertical wind, D-A-CH in Leipzig,
Ideal conditions (red lines indicate 500/1000/1500 m)
Red area indicate blind zone up to 85 m due to pulse length of Lidar system (320 ns)

## Slide 18: Pulsed Doppler Lidar Wind Ranger BL, 24.08.2022, 13:44 UTC, 10 s Averages
Velocity
Direction
SNR
U-Comp.
V-Comp.
W-Comp.
white_ estimated BL red available Lidar data
all 20 minutes

## Slide 19: Wind Ranger 100/200 + Wind Ranger BL = Wind Ranger XM
Wind Ranger
100/200
VAD scans at 80° elevation at a rate of 0.5 … 2 rps
Sampe rate 100 Hz
No 180° ambiguity, no blind zero zone
Up to 20 measuring heights available (typ. 6 … 10)
Determination of the effective measuring range
Measuring heights 7 … 200 m
Narrow scan angle 10 °, in favor of vertical wind
Wind Ranger
BL
+
VAD scans at 80° elevation at a rate of 0.1 rps
Pulse rate 10 kHz, pulse length 320 ns, i.e. 100 m
Up to 100 measuring heights available
Measuring heights 80 … 2500 m
Narrow scan angle 10 °, in favor of vertical wind
=
Wind Ranger
XM
Combined staggered mode of both techniques
Typ. 4/8 minutes low mode, 1/2 minutes high mode

## Slide 20: Capability of Wind Ranger XM Data samples of combined modes
Pulse Mode
1500 m … 72 m
FMCW Mode
200 m … 10 m
Height (m)
Height (m)
Wind Speed (m/s)
15
10
5
0
11.08 12:00
12.08 00:00
12.08 12:00
13.08 00:00
13.08 12:00
14.08 00:00
14.08 12:00
15
10
5
0
Wind Speed (m/s)
10
Pls. regard the different color scaling

## Slide 21: Capability of Wind Ranger XM Data samples of combined modes, WR BL data
10
(850 … 189 m)

## Slide 22: Capability of Wind Ranger XM Data samples of combined modes, WR 200 data
10
(200 … 10 m)
(850 … 189 m)

## Slide 23: Capability of Wind Ranger XM Data samples of combined modes, Low Level Jet 1
10
Time development of with
,
East-West-component at 230 m height,
ast-West-component at 850 m height,
South-North-component at 230 m height,
South-North-component at 850 m height.
10
5
U-U
g
(m/s)
0
-5
10
V-V
g
(m/s)
5
0
-10
-5
-10
11-12.08.2022

## Slide 24: Capability of Wind Ranger XM Data samples of combined modes, Low Level Jet 2
10
10
5
U-U
g
(m/s)
0
10
5
-5
V-V
g
(m/s)
0
-5
-10
-10
12-13.08.2022

## Slide 25: Capability of Wind Ranger XM Data samples of combined modes, Low Level Jet 3
10
-10
-5
0
5
10
U-U
g
(m/s)
-10
-5
0
5
10
V-V
g
(m/s)
13-14.08.2022

## Slide 26: Capability of Wind Ranger XM Standard deviation of vertical wind, 200 m

## Slide 27: Capability of Wind Ranger XM Standard deviation of vertical wind, 160 m

## Slide 28: Capability of Wind Ranger XM Standard deviation of vertical wind, 100 m

## Slide 29: Capability of Wind Ranger XM Standard deviation of vertical wind, 60 m

## Slide 30: Capability of Wind Ranger XM Standard deviation of vertical wind, 40 m

## Slide 31: Capability of Wind Ranger XM Standard deviation of vertical wind, 20 m

## Slide 32: Capability of Wind Ranger XM Standard deviation of vertical wind, 10 m
Decrease in 10 m height because measuring site is surrounded by 10 m high buildíngs !

## Slide 33: Conclusions and Outlook
Compact and light weight Lidar systems can combine low range/ long range measurements
Fine resolution in low ranges, 8 cm in 10 m height
Lowest heights down to 10 m
In long range mode (100 m pulse length) measurements covering the upper limit of observed PBL are possible
Turbulence measurements look quite promising for WR 100/200, WR XM still open
Verifications with masts and short range/long range lidars are next topics
Monitoring the „whole“ PBL in good weather conditions seems feasible

## Slide 34: Thank you for your attention